# CHIVA Streaming — Test Set 4 Report (Shunt Types I–VI)

Date: 2026-06-19  11:00:57    |    API: http://localhost:7861    |    Scenarios: 6

# Overall Summary

| Scenario | Description | Passed | Failed | Timeouts |
| --- | --- | --- | --- | --- |
| type1 | Type I (N1→N2→N1): SFJ incompetent entry (EP N1→N2); GSV trunk reflux; no tributary escape; re-entry from trunk directly via calf perforator (RP N2→N1). Rule 6 complete (max_visited ≥ 0.44). | 22 | 0 | 0 |
| type2 | Type II (N2→N3→N1): Both junctions competent; saphenous trunk is reflux source; trunk overflows to tributary (EP N2→N3); tributary re-enters deep (RP N3→N1). No EP N1→N2; no RP N2→N1. Rule 7 complete. | 22 | 0 | 0 |
| type3 | Type III (N1→N2→N3→N1): SFJ entry (EP N1→N2); ALL volume conducted via tributary (EP N2→N3 before RP N2→N1); RP N3→N1 + RP N2→N1 → Maneuver; elim='No Reflux' → Type III complete (trunk was conduit). | 23 | 0 | 0 |
| type4 | Type IV (N1→N3→N2→N1): Both junctions competent; pelvic/perf exits into tributary (EP N1→N3); tributary fills trunk; trunk re-enters deep (RP N2→N1). Rule 4 complete immediately (no max_visited gate). | 23 | 0 | 0 |
| type5 | Type V (N1→N3→N2→N3→N1): Source perforator into tributary (EP N1→N3); tributary drains into trunk (RP N3→N2); trunk escapes to second tributary (EP N2→N3); second tributary re-enters deep (RP N3→N1). Rule 5b complete. | 24 | 0 | 0 |
| type6 | Type VI (N1→N3→N1): Both junctions competent; perforator exits into tributary (EP N1→N3); PURELY superficial N3 path; re-entry perforator (RP N3→N1). No trunk (N2) involvement. Rule 5 complete. | 22 | 0 | 0 |

OVERALL: ALL PASS

# Scenario: TYPE1

Type I (N1→N2→N1): SFJ incompetent entry (EP N1→N2); GSV trunk reflux; no tributary escape; re-entry from trunk directly via calf perforator (RP N2→N1). Rule 6 complete (max_visited ≥ 0.44).

Checked: 22   |   Passed: 22   |   Failed: 0   |   Timeouts: 0

| # | Surgeon Action | Expected | LLM Guidance | Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: ts4_type1) | session_ready event | -- | -- | SKIP |
| 2 | Move probe → SFJ posY=0.06 anterior-medial right leg | Action = "move" Guidance contains any of: SFJ, groin, femoral, junction | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 3 | Move probe → SFJ posY=0.09 anterior-medial right leg | Action = "move" | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 4 | Mark: EP N1→N2 posY=0.07 SFJ right leg | Action = "move" | Move distally along medial thigh | move | PASS |
| 5 | Move probe → UPPER_THIGH posY=0.13 medial right leg | Action = "move" | Track GSV distally along medial thigh | move | PASS |
| 6 | Move probe → UPPER_THIGH posY=0.19 medial right leg | Action = "move" | Move distally along medial thigh | move | PASS |
| 7 | Move probe → HUNTERIAN posY=0.26 medial right leg | Action = "move" | Track GSV distally along medial thigh | move | PASS |
| 8 | Move probe → HUNTERIAN posY=0.32 medial right leg | Action = "move" | Move distally toward knee | move | PASS |
| 9 | Move probe → HUNTERIAN posY=0.38 medial right leg | Action = "move" | Move distally toward popliteal fossa | move | PASS |
| 10 | Move probe → POPLITEAL posY=0.47 posterior right leg | Action = "move" Guidance contains any of: popliteal, SPJ, posterior | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 11 | Move probe → POPLITEAL posY=0.53 posterior right leg | Action = "move" | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 12 | Move probe → CALF posY=0.60 medial right leg | Action = "move" | Track GSV distally along medial calf | move | PASS |
| 13 | Move probe → CALF posY=0.66 medial right leg | Action = "move" | Follow GSV distally on medial calf | move | PASS |
| 14 | Move probe → CALF posY=0.72 medial right leg | Action = "move" | Track medial calf distally | move | PASS |
| 15 | Move probe → ANKLE posY=0.85 medial right leg | Action = "move" | Scan medial ankle | move | PASS |
| 16 | Move probe → CALF posY=0.79 medial right leg | Action = "move" | Track GSV distally on medial calf | move | PASS |
| 17 | Move probe → CALF posY=0.73 medial right leg | Action = "move" | Follow medial calf distally | move | PASS |
| 18 | Move probe → CALF posY=0.67 medial right leg | Action = "move" | Track GSV on medial calf | move | PASS |
| 19 | Move probe → HUNTERIAN posY=0.30 medial right leg | Action = "move" | Track GSV distally along medial thigh | move | PASS |
| 20 | Move probe → HUNTERIAN posY=0.24 medial right leg | Action = "move" | Move distally along medial thigh | move | PASS |
| 21 | Mark: RP N2→N1 posY=0.64 CALF right leg | Action = "complete" Guidance contains any of: circuit, classification, complete, mapped | Circuit mapped — sufficient findings for classification | complete | PASS |
| 22 | Move probe → POPLITEAL posY=0.50 posterior right leg | Action = "complete" Guidance contains any of: circuit, classification | Circuit mapped — sufficient findings for classification | complete | PASS |
| 23 | Move probe → SFJ posY=0.07 anterior-medial right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |

# Scenario: TYPE2

Type II (N2→N3→N1): Both junctions competent; saphenous trunk is reflux source; trunk overflows to tributary (EP N2→N3); tributary re-enters deep (RP N3→N1). No EP N1→N2; no RP N2→N1. Rule 7 complete.

Checked: 22   |   Passed: 22   |   Failed: 0   |   Timeouts: 0

| # | Surgeon Action | Expected | LLM Guidance | Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: ts4_type2) | session_ready event | -- | -- | SKIP |
| 2 | Move probe → SFJ posY=0.06 anterior-medial right leg | Action = "move" Guidance contains any of: SFJ, groin, femoral, junction | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 3 | Move probe → UPPER_THIGH posY=0.12 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 4 | Move probe → UPPER_THIGH posY=0.18 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 5 | Move probe → HUNTERIAN posY=0.24 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 6 | Move probe → HUNTERIAN posY=0.30 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 7 | Move probe → HUNTERIAN posY=0.36 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 8 | Move probe → POPLITEAL posY=0.47 posterior right leg | Action = "move" Guidance contains any of: popliteal, SPJ, posterior | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 9 | Move probe → CALF posY=0.55 posterior right leg | Action = "move" | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 10 | Move probe → CALF posY=0.61 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 11 | Mark: EP N2→N3 posY=0.29 HUNTERIAN right leg | Action = "move" Must NOT = "complete" | Move distally along medial thigh | move | PASS |
| 12 | Move probe → DISTAL_THIGH posY=0.40 medial right leg | Action = "move" | Move distally along medial thigh | move | PASS |
| 13 | Move probe → POPLITEAL posY=0.49 posterior right leg | Action = "move" | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 14 | Move probe → CALF posY=0.56 posterior right leg | Action = "move" | Move distally along posterior calf | move | PASS |
| 15 | Move probe → CALF posY=0.63 posterior right leg | Action = "move" | Move distally along posterior calf | move | PASS |
| 16 | Move probe → CALF posY=0.69 posterior right leg | Action = "move" | Move distally along posterior calf | move | PASS |
| 17 | Move probe → CALF posY=0.75 lateral right leg | Action = "move" | Move distally along lateral calf | move | PASS |
| 18 | Move probe → ANKLE posY=0.85 lateral right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 19 | Move probe → CALF posY=0.78 lateral right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 20 | Move probe → CALF posY=0.71 lateral right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 21 | Mark: RP N3→N1 posY=0.68 CALF right leg | Action = "complete" Guidance contains any of: circuit, classification, complete, mapped | Circuit mapped — sufficient findings for classification | complete | PASS |
| 22 | Move probe → HUNTERIAN posY=0.28 medial right leg | Action = "complete" Guidance contains any of: circuit, classification | Circuit mapped — sufficient findings for classification | complete | PASS |
| 23 | Move probe → SFJ posY=0.07 anterior-medial right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |

# Scenario: TYPE3

Type III (N1→N2→N3→N1): SFJ entry (EP N1→N2); ALL volume conducted via tributary (EP N2→N3 before RP N2→N1); RP N3→N1 + RP N2→N1 → Maneuver; elim='No Reflux' → Type III complete (trunk was conduit).

Checked: 23   |   Passed: 23   |   Failed: 0   |   Timeouts: 0

| # | Surgeon Action | Expected | LLM Guidance | Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: ts4_type3) | session_ready event | -- | -- | SKIP |
| 2 | Move probe → SFJ posY=0.06 anterior-medial right leg | Action = "move" Guidance contains any of: SFJ, groin, femoral, junction | Scan transversely at groin crease | move | PASS |
| 3 | Mark: EP N1→N2 posY=0.07 SFJ right leg | Action = "move" | Move distally along medial thigh | move | PASS |
| 4 | Move probe → UPPER_THIGH posY=0.13 medial right leg | Action = "move" | Track GSV distally on medial thigh | move | PASS |
| 5 | Move probe → UPPER_THIGH posY=0.20 medial right leg | Action = "move" | Move distally along medial thigh | move | PASS |
| 6 | Move probe → HUNTERIAN posY=0.26 medial right leg | Action = "move" | Continue along medial thigh | move | PASS |
| 7 | Mark: EP N2→N3 posY=0.28 HUNTERIAN right leg | Action = "move" Must NOT = "complete" | Follow tributary distally on medial thigh | move | PASS |
| 8 | Move probe → POPLITEAL posY=0.47 posterior right leg | Action = "move" Guidance contains any of: popliteal, SPJ, posterior | Scan posteriorly at popliteal fossa | move | PASS |
| 9 | Move probe → POPLITEAL posY=0.53 posterior right leg | Action = "move" | Scan posteriorly at popliteal fossa | move | PASS |
| 10 | Move probe → CALF posY=0.60 posterior right leg | Action = "move" | Track N3 tributary distally on posterior calf | move | PASS |
| 11 | Move probe → CALF posY=0.66 medial right leg | Action = "move" | Follow tributary on medial calf | move | PASS |
| 12 | Move probe → CALF posY=0.72 medial right leg | Action = "move" | Track medial calf distally | move | PASS |
| 13 | Move probe → ANKLE posY=0.85 medial right leg | Action = "move" | Follow N3 tributary on medial ankle | move | PASS |
| 14 | Move probe → CALF posY=0.79 medial right leg | Action = "move" | Track medial calf distally | move | PASS |
| 15 | Move probe → CALF posY=0.73 medial right leg | Action = "move" | Follow tributary on medial calf | move | PASS |
| 16 | Mark: RP N3→N1 posY=0.69 CALF right leg | Action = "move" Must NOT = "maneuver" | Continue scanning distally to locate anatomical junction | move | PASS |
| 17 | Move probe → HUNTERIAN posY=0.32 medial right leg | Action = "move" | Continue scanning distally to locate anatomical junction | move | PASS |
| 18 | Move probe → HUNTERIAN posY=0.27 medial right leg | Action = "move" | Continue scanning distally to locate anatomical junction | move | PASS |
| 19 | Mark: RP N2→N1 posY=0.25 HUNTERIAN right leg | Action = "maneuver" Guidance contains any of: compress, tributary | Compress tributary — record whether GSV Doppler changes | maneuver | PASS |
| 20 | Move probe → CALF posY=0.64 medial right leg | Action = "maneuver" | Compress tributary — record whether GSV Doppler changes | maneuver | PASS |
| 21 | Move probe → HUNTERIAN posY=0.30 medial right leg | Action = "maneuver" | Compress tributary — record whether GSV Doppler changes | maneuver | PASS |
| 22 | Mark: EP N2→N3 posY=0.28 HUNTERIAN right leg Elim: No Reflux | Action = "complete" Guidance contains any of: circuit, classification, complete, mapped | Circuit mapped — sufficient findings for classification | complete | PASS |
| 23 | Move probe → CALF posY=0.65 medial right leg | Action = "complete" Guidance contains any of: circuit, classification | Circuit mapped — sufficient findings for classification | complete | PASS |
| 24 | Move probe → SFJ posY=0.07 anterior-medial right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |

# Scenario: TYPE4

Type IV (N1→N3→N2→N1): Both junctions competent; pelvic/perf exits into tributary (EP N1→N3); tributary fills trunk; trunk re-enters deep (RP N2→N1). Rule 4 complete immediately (no max_visited gate).

Checked: 23   |   Passed: 23   |   Failed: 0   |   Timeouts: 0

| # | Surgeon Action | Expected | LLM Guidance | Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: ts4_type4) | session_ready event | -- | -- | SKIP |
| 2 | Move probe → SFJ posY=0.06 anterior-medial right leg | Action = "move" Guidance contains any of: SFJ, groin, femoral, junction | Scan transversely at groin crease | move | PASS |
| 3 | Move probe → POPLITEAL posY=0.47 posterior right leg | Action = "move" Guidance contains any of: popliteal, SPJ, posterior | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 4 | Move probe → CALF posY=0.55 posterior right leg | Action = "move" | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 5 | Move probe → CALF posY=0.61 medial right leg | Action = "move" | Move distally along medial calf | move | PASS |
| 6 | Move probe → CALF posY=0.67 medial right leg | Action = "move" | Track medial calf distally | move | PASS |
| 7 | Move probe → CALF posY=0.73 medial right leg | Action = "move" | Follow GSV distally on medial calf | move | PASS |
| 8 | Move probe → ANKLE posY=0.85 medial right leg | Action = "move" | Move proximally toward calf | move | PASS |
| 9 | Move probe → ANKLE posY=0.91 lateral right leg | Action = "move" | Move proximally toward calf | move | PASS |
| 10 | Move probe → ANKLE posY=0.85 medial right leg | Action = "move" | Move proximally toward calf | move | PASS |
| 11 | Move probe → CALF posY=0.78 medial right leg | Action = "move" | Move proximally toward popliteal fossa | move | PASS |
| 12 | Mark: EP N1→N3 posY=0.71 CALF right leg | Action = "move" Must NOT = "complete" | Track N3 tributary distally on medial calf | move | PASS |
| 13 | Move probe → POPLITEAL posY=0.50 posterior right leg | Action = "move" | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 14 | Move probe → DISTAL_THIGH posY=0.42 medial right leg | Action = "move" | Move distally along medial thigh | move | PASS |
| 15 | Move probe → HUNTERIAN posY=0.36 medial right leg | Action = "move" | Move distally along medial thigh | move | PASS |
| 16 | Move probe → HUNTERIAN posY=0.30 medial right leg | Action = "move" | Move distally along medial thigh | move | PASS |
| 17 | Move probe → HUNTERIAN posY=0.24 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 18 | Move probe → UPPER_THIGH posY=0.17 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 19 | Move probe → UPPER_THIGH posY=0.13 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 20 | Move probe → HUNTERIAN posY=0.20 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 21 | Move probe → HUNTERIAN posY=0.27 medial right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 22 | Mark: RP N2→N1 posY=0.25 HUNTERIAN right leg | Action = "complete" Guidance contains any of: circuit, classification, complete, mapped | Circuit mapped — sufficient findings for classification | complete | PASS |
| 23 | Move probe → CALF posY=0.65 medial right leg | Action = "complete" Guidance contains any of: circuit, classification | Circuit mapped — sufficient findings for classification | complete | PASS |
| 24 | Move probe → SFJ posY=0.07 anterior-medial right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |

# Scenario: TYPE5

Type V (N1→N3→N2→N3→N1): Source perforator into tributary (EP N1→N3); tributary drains into trunk (RP N3→N2); trunk escapes to second tributary (EP N2→N3); second tributary re-enters deep (RP N3→N1). Rule 5b complete.

Checked: 24   |   Passed: 24   |   Failed: 0   |   Timeouts: 0

| # | Surgeon Action | Expected | LLM Guidance | Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: ts4_type5) | session_ready event | -- | -- | SKIP |
| 2 | Move probe → SFJ posY=0.06 anterior-medial right leg | Action = "move" Guidance contains any of: SFJ, groin, femoral, junction | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 3 | Move probe → POPLITEAL posY=0.47 posterior right leg | Action = "move" Guidance contains any of: popliteal, SPJ, posterior | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 4 | Move probe → CALF posY=0.55 posterior right leg | Action = "move" | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 5 | Move probe → CALF posY=0.61 medial right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 6 | Move probe → CALF posY=0.67 medial right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 7 | Mark: EP N1→N3 posY=0.64 CALF right leg | Action = "move" Must NOT = "complete" | Track medial calf distally following N3 tributary | move | PASS |
| 8 | Move probe → DISTAL_THIGH posY=0.39 medial right leg | Action = "move" | Move distally along medial thigh toward calf | move | PASS |
| 9 | Move probe → HUNTERIAN posY=0.33 medial right leg | Action = "move" | Move distally along medial thigh toward calf | move | PASS |
| 10 | Move probe → HUNTERIAN posY=0.27 medial right leg | Action = "move" | Move distally along medial thigh toward calf | move | PASS |
| 11 | Move probe → HUNTERIAN posY=0.21 medial right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 12 | Move probe → HUNTERIAN posY=0.28 medial right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 13 | Mark: RP N3→N2 posY=0.30 HUNTERIAN right leg | Action = "move" Must NOT = "complete" | Move distally along medial thigh toward calf | move | PASS |
| 14 | Move probe → CALF posY=0.57 medial right leg | Action = "move" | Track medial calf distally following N3 tributary | move | PASS |
| 15 | Move probe → CALF posY=0.63 medial right leg | Action = "move" | Track medial calf distally following N3 tributary | move | PASS |
| 16 | Move probe → CALF posY=0.69 medial right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 17 | Move probe → CALF posY=0.75 medial right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 18 | Mark: EP N2→N3 posY=0.72 CALF right leg | Action = "move" Must NOT = "complete" | Track medial calf distally following N3 tributary | move | PASS |
| 19 | Move probe → CALF posY=0.78 medial right leg | Action = "move" | Track medial calf distally following N3 tributary | move | PASS |
| 20 | Move probe → ANKLE posY=0.86 medial right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 21 | Move probe → ANKLE posY=0.92 medial right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 22 | Move probe → ANKLE posY=0.86 medial right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 23 | Mark: RP N3→N1 posY=0.89 ANKLE right leg | Action = "complete" Guidance contains any of: circuit, classification, complete, mapped | Circuit mapped — sufficient findings for classification | complete | PASS |
| 24 | Move probe → CALF posY=0.65 medial right leg | Action = "complete" Guidance contains any of: circuit, classification | Circuit mapped — sufficient findings for classification | complete | PASS |
| 25 | Move probe → HUNTERIAN posY=0.29 medial right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |

# Scenario: TYPE6

Type VI (N1→N3→N1): Both junctions competent; perforator exits into tributary (EP N1→N3); PURELY superficial N3 path; re-entry perforator (RP N3→N1). No trunk (N2) involvement. Rule 5 complete.

Checked: 22   |   Passed: 22   |   Failed: 0   |   Timeouts: 0

| # | Surgeon Action | Expected | LLM Guidance | Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: ts4_type6) | session_ready event | -- | -- | SKIP |
| 2 | Move probe → SFJ posY=0.06 anterior-medial right leg | Action = "move" Guidance contains any of: SFJ, groin, femoral, junction | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 3 | Move probe → POPLITEAL posY=0.47 posterior right leg | Action = "move" Guidance contains any of: popliteal, SPJ, posterior | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 4 | Move probe → CALF posY=0.55 posterior right leg | Action = "move" | Scan posteriorly at popliteal fossa to assess SPJ | move | PASS |
| 5 | Move probe → CALF posY=0.61 medial right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 6 | Move probe → CALF posY=0.64 lateral right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 7 | Move probe → CALF posY=0.70 lateral right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 8 | Move probe → CALF posY=0.76 lateral right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 9 | Move probe → ANKLE posY=0.85 lateral right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 10 | Move probe → ANKLE posY=0.91 lateral right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 11 | Move probe → ANKLE posY=0.85 lateral right leg | Action = "move" | Move proximally toward groin to assess femoral junction | move | PASS |
| 12 | Mark: EP N1→N3 posY=0.87 ANKLE right leg | Action = "move" Must NOT = "complete" | Track N3 tributary distally on lateral ankle | move | PASS |
| 13 | Move probe → CALF posY=0.80 lateral right leg | Action = "move" | Move distally along lateral calf | move | PASS |
| 14 | Move probe → CALF posY=0.74 lateral right leg | Action = "move" | Move distally along lateral calf | move | PASS |
| 15 | Move probe → CALF posY=0.68 lateral right leg | Action = "move" | Move distally along lateral calf | move | PASS |
| 16 | Move probe → CALF posY=0.62 posterior right leg | Action = "move" | Move distally along posterior calf | move | PASS |
| 17 | Move probe → CALF posY=0.66 medial right leg | Action = "move" | Move distally along medial calf | move | PASS |
| 18 | Move probe → CALF posY=0.73 lateral right leg | Action = "move" | Move distally along lateral calf | move | PASS |
| 19 | Move probe → CALF posY=0.67 lateral right leg | Action = "move" | Move distally along lateral calf | move | PASS |
| 20 | Mark: RP N3→N1 posY=0.71 CALF right leg | Action = "complete" Guidance contains any of: circuit, classification, complete, mapped | Circuit mapped — sufficient findings for classification | complete | PASS |
| 21 | Move probe → CALF posY=0.67 lateral right leg | Action = "complete" Guidance contains any of: circuit, classification | Circuit mapped — sufficient findings for classification | complete | PASS |
| 22 | Move probe → POPLITEAL posY=0.48 posterior right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |
| 23 | Move probe → SFJ posY=0.07 anterior-medial right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |