# CHIVA Streaming Guidance — Extended Scenario Test Report (v2)

Date: 2026-06-25  19:30:27    |    API: http://localhost:7861    |    Scenarios: 1
Literature: Franceschi & Zamboni 2009 · Gianesini 2014 · Delfrate 2023 · Adler 2022

# Overall Summary

| Scenario | Description | Passed | Failed | Timeouts |
| --- | --- | --- | --- | --- |
| rule6_max_visited_gate | Rule 6 max_visited gate regression: EP N1→N2 SFJ + RP N2→N1 at posY=0.15 (old_max<0.48) → forbidden complete until probe covers SPJ zone. Complete fires only after posY≥0.48. Source: run_scenarios.py gate spec. 12 steps. | 8 | 3 | 3 |

OVERALL: FAILURES PRESENT

# Scenario: RULE6_MAX_VISITED_GATE

Rule 6 max_visited gate regression: EP N1→N2 SFJ + RP N2→N1 at posY=0.15 (old_max<0.48) → forbidden complete until probe covers SPJ zone.  Complete fires only after posY≥0.48. Source: run_scenarios.py gate spec.  12 steps.

Checked: 11   |   Passed: 8   |   Failed: 3   |   Timeouts: 3

| # | Surgeon Action | Expected | LLM Guidance | Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: s14_gate) | session_ready event | -- | -- | SKIP |
| 2 | Move → SFJ posY=0.05 anterior-medial right leg | Action = "move" Guidance contains any of: SFJ, groin, femoral, junction | Move probe distally to SFJ | move | PASS |
| 3 | Move → SFJ posY=0.07 anterior-medial right leg | Action = "move" | -- | -- | TIMEOUT |
| 4 | Mark: EP N1→N2 posY=0.06 SFJ right leg | Action = "move" | Move probe distally to mid-thigh GSV | move | PASS |
| 5 | Move → Upper Thigh posY=0.12 medial right leg | Action = "move" | Move probe distally to GSV | move | PASS |
| 6 | Move → Upper Thigh posY=0.16 medial right leg | Action = "move" | -- | -- | TIMEOUT |
| 7 | Mark: RP N2→N1 posY=0.15 Upper Thigh right leg | Action = "move" Must NOT = "complete" | Move probe distally to mid-thigh GSV | move | PASS |
| 8 | Move → Dodd posY=0.24 medial right leg | Action = "move" Must NOT = "complete" | Move probe distally to mid-thigh GSV | move | PASS |
| 9 | Move → Dodd posY=0.30 medial right leg | Action = "move" Must NOT = "complete" | Move probe distally to mid-thigh GSV | move | PASS |
| 10 | Move → Hunterian posY=0.36 medial right leg | Action = "move" Must NOT = "complete" | Move probe distally to Hunterian perforator | move | PASS |
| 11 | Move → Hunterian posY=0.46 medial right leg | Action = "move" Must NOT = "complete" | Move probe distally to Hunterian perforator | move | PASS |
| 12 | Move → SPJ posY=0.49 posterior right leg | Action = "move" Guidance contains any of: popliteal, SPJ, posterior, junction | -- | -- | TIMEOUT |
| 13 | Move → SPJ posY=0.53 posterior right leg | Action = "complete" Guidance contains any of: circuit, classification, complete, mapped | Move probe distally to mid-thigh GSV | move | FAIL |
| x Action: got "move", expected "complete" x Guidance missing any of: circuit, classification, complete, mapped |  |  |  |  |  |
| 14 | Move → Hunterian posY=0.40 medial right leg | Action = "complete" Guidance contains any of: circuit, classification | Move probe distally to SFJ | move | FAIL |
| x Action: got "move", expected "complete" x Guidance missing any of: circuit, classification |  |  |  |  |  |
| 15 | Move → SFJ posY=0.06 anterior-medial right leg | Action = "complete" | Move probe distally to SFJ | move | FAIL |
| x Action: got "move", expected "complete" |  |  |  |  |  |