# Cygnus Med – Task-2 Active Guidance System

# Clinical Scenario Test Report: Type 1 + 2 Shunt

Test date:  2026-07-10 14:47
Source video:  Media - Shunt type 1.mp4  (3 min 4 s, Mindray M9CV, 8 MHz linear)
Patient:  Standing, right leg
Diagnosis:  Type 1 + 2 combined shunt (CHIVA classification)
Backend:  http://127.0.0.1:7861

## 1. Examination Narrative (from video)

0–40 s   SFJ Assessment: Surgeon places probe at groin crease. B-mode: CFV + GSV adjacent ovals. Colour Doppler activated — large BLUE fill (CFV antegrade), then on calf release: large solid RED blob fills GSV = retrograde flow = SFJ incompetence. TYPE 1 ESCAPE POINT confirmed.

40–80 s  GSV Thigh Tracking: Probe swept distal along anteromedial thigh. Colour Doppler: solid red fill throughout GSV from upper thigh → Hunterian canal → Dodd zone = retrograde. PW Doppler at Hunterian (SVD 16.6 mm, Angle −60°): downward spectral trace confirms retrograde GSV flow. TYPE 1 SHUNT PATH fully mapped.

80–130 s Lower Leg Scan: Probe moved to calf. B-mode: small GSV oval in fascial compartment at each level from upper to ankle. Multiple vessels visible during transit at knee (popliteal structures). GSV reflux column mapped to ankle level.

130–160 s SPJ Assessment: Probe at popliteal fossa (posterior knee). B-mode: TWO large adjacent ovals = popliteal vein (N1) + SSV (N2). Colour Doppler: RED signal at their interface = SSV reflux into popliteal vein. PW gate set (Angle 60°). TYPE 2 ESCAPE POINT confirmed — SPJ incompetence.

160–183 s Elimination Test: Probe returns to SFJ. Colour Doppler: bidirectional (RED+BLUE). PW Doppler trace captured: upward peak (antegrade, calf squeeze) + downward deflection (retrograde, release). POSITIVE ELIMINATION TEST — SFJ confirmed as the single escape point for the Type 1 shunt. FINAL DIAGNOSIS: Type 1 + 2 shunt.

## 2. Summary Scoreboard

| Metric | Score | Assessment |
| --- | --- | --- |
| Region localisation accuracy | 11/11 | PASS |
| LLM guidance generated (non-error) | 11/11 | PASS |
| Guidance clinically CORRECT | 3/11 | CORRECT=3 PARTIAL=4 WRONG=4 NO_RESP=0 |
| LLM prompt contains clinical context | 11/11 | PASS |
| Vein frame endpoint returns image | 11/11 | PASS |
| Wrong-region detection | 4/4 | PASS |

## 3. Step-by-Step Clinical Assessment

| Time (video) | Phase | What Surgeon Did (Video Finding) | Probe Position Sent to System | Region Detected | Guidance Text Generated by System | Expected Next Step (Ground Truth) | Guidance Score | Prompt Context OK? | Vein Frame | Wrong-Region Detected? | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0–5 s | Ph1 SFJ Assessment | Surgeon places probe at groin crease, anterior surface. B-mode: small anechoic oval (GSV) visible beside larger CFV. Finding: SFJ identified. | Groin crease, anterior (dist=0.05) | SFJ (✓ exp: SFJ) | Look for N1 alongside N2 at SFJ junction level. | Switch Colour Doppler; confirm reflux at SFJ junction (GSV junction with CFV). | PARTIAL | YES | YES GSV proximal trunk | YES ✓ Probe is scanning posterior anatomy (SPJ) but the protocol r | Video 0 s: B-mode only, probe at inner groin. CFV + GSV visible as adjacent anechoic ovals. |
| 10–30 s | Ph1 SFJ Assessment | Colour Doppler activated. Large BLUE fill (CFV antegrade), then mixed red/blue at SFJ junction. At ~20 s: large solid RED blob filling GSV — reflux wave on calf-release. Finding: Significant SFJ incompetence confirmed (Type 1 escape point). | SFJ, with Colour Doppler active (dist=0.07) | SFJ (✓ exp: SFJ) | Move probe to upper thigh, posY≈0.10, to assess GSV trunk. | Confirm SFJ junction anatomy on B-mode, then track GSV distally through thigh to map reflux column. | WRONG | YES | YES GSV proximal trunk | N/A | Video 10–30 s: Colour Doppler shows solid red on calf release = retrograde flow = SFJ reflux. Type 1 shunt escape confirmed. |
| 40–50 s | Ph2 GSV Thigh Tracking | Probe moved to upper thigh (anterior medial). B-mode: large anechoic GSV oval in fascial compartment. Colour Doppler: solid red fill = retrograde flow in GSV trunk. Finding: GSV reflux confirmed in upper thigh (Hunterian canal zone). | Upper thigh, anterior (dist=0.30) — Hunterian proximal | GSV-THI (✓ exp: GSV-THI) | Move probe to mid-thigh, posY≈0.25, to assess Hunterian junction. | Continue distally along thigh, confirm continuous GSV reflux column down to calf; look for N3 tributary escape point. | PARTIAL | YES | YES GSV thigh | YES ✓ Probe detected at SFJ (Saphenofemoral Junction) but expected | Video 40–50 s: Probe at mid-thigh, solid red Colour Doppler fill = large-calibre retrograde GSV. Type 1 pathway confirmed in thigh. |
| 50–65 s | Ph2 GSV Thigh Tracking | PW Doppler activated at mid-thigh. Gate placed on GSV. Spectral trace shows downward waveform below baseline (SVD≈16.6 mm, Angle −60°). Finding: Spectral confirmation of retrograde GSV flow at thigh level. | Mid-thigh / Hunterian canal, anterior (dist=0.48) | GSV-THI (✓ exp: GSV-THI) | Move probe to popliteal fossa, posY≈0.45, to assess SPJ junction. | Move probe further distally along thigh into distal GSV / Dodd zone; continue mapping reflux column toward calf. | WRONG | YES | YES GSV distal trunk | N/A | Video 55–65 s: PW Doppler trace clearly below baseline = retrograde. SVD 16.6 mm confirms significant GSV trunk diameter. This is the CHIVA RP N2→N1 confirmation step. |
| 65–80 s | Ph2 GSV Thigh Tracking | Probe moved to distal thigh (Dodd perforator zone). B-mode: single GSV oval still visible in fascial compartment. Finding: GSV reflux column extends to distal thigh. | Distal thigh / Dodd zone, anterior (dist=0.65) | GSV-THI (✓ exp: GSV-THI) | Move probe to mid-calf, posY≈0.60, to assess GSV-tributary junction. | Move probe to calf, track GSV into fascial compartment; map reflux extent to calf/ankle. | CORRECT | YES | YES GSV distal trunk | N/A | Video 65–80 s: Probe moves from knee-level back to mid-thigh. Fascia clearly visible, GSV in compartment throughout. Dodd perforator zone is CHIVA Type 2 sub-type boundary. |
| 80–100 s | Ph3 Lower Leg Scan | Probe moved to upper calf (just below knee), anterior medial. B-mode: small anechoic GSV oval visible in fascial compartment. Finding: GSV continues into calf with same N2 appearance. | Upper calf, anterior medial (seg=1, dist=0.15) | GSV-CAL (✓ exp: GSV-CAL) | Move probe to correct calf position, posY≈0.60, to assess GSV-tributary. | Scan distally along calf to ankle, map full extent of GSV reflux column; look for N3 tributary escape at calf level. | CORRECT | YES | YES GSV distal trunk | YES ✓ Probe is scanning posterior anatomy (SPJ) but the protocol r | Video 80–100 s: Probe at upper-mid calf. Camera shows probe at shin level. Small single GSV oval on B-mode. |
| 100–120 s | Ph3 Lower Leg Scan | Probe swept to lower calf / ankle area. B-mode: small GSV oval at ankle level. Multiple vessels visible at popliteal during transit. Finding: GSV reflux maps to ankle, large cluster at knee = popliteal structures. | Lower calf / ankle, anterior medial (seg=1, dist=0.75) | GSV-CAL (✓ exp: GSV-CAL) | Move probe to ankle, posY≈0.85, to assess GSV termination. | Reassess posterior knee (popliteal fossa) to evaluate SPJ and check for Type 2 shunt via SSV. | PARTIAL | YES | YES GSV distal trunk | N/A | Video 100–120 s: Probe clearly at ankle. Camera shows probe at foot level. GSV very small calibre at this level. |
| 130–145 s | Ph4 SPJ Assessment | Probe placed at popliteal fossa (posterior knee). B-mode: two large adjacent anechoic ovals — popliteal vein (N1, larger) and SSV (N2, smaller). Colour Doppler: small RED signal at SSV/popliteal interface = reflux. Finding: SPJ incompetence confirmed — SSV reflux into popliteal ( | Popliteal fossa, posterior (dist=0.87, back of thigh) | SPJ (✓ exp: SPJ) | Move probe to groin crease, posY≈0.06, to assess SFJ junction. | Confirm SSV reflux with PW Doppler at SPJ; then return to SFJ for elimination test. | WRONG | YES | YES PV (popliteal vein) | YES ✓ Probe is scanning anterior anatomy (GSV-CAL) but the protoco | Video 130 s: Two large ovals at popliteal. Red signal at interface. Video 140 s: PW gate positioned, Angle 60°. This is the Type 2 shunt confirmation — SPJ incompetence. |
| 145–160 s | Ph4 SPJ Assessment | B-mode at SPJ shows 4 anechoic ovals (popliteal vein, SSV, tributaries). PW Doppler gate set at Angle 60°. Finding: Full SPJ anatomy mapped, multiple vessels at junction. | Posterior upper calf / SPJ from below (seg=1, dist=0.03) | SPJ (✓ exp: SPJ) | Adjust probe to anterior thigh, posY≈0.06, to confirm SFJ junction. | Return probe to groin (SFJ). Perform elimination test with calf compression under PW Doppler. | PARTIAL | YES | YES SSV (small saphenous) | N/A | Video 150 s: 4 vessels visible at SPJ — popliteal v., SSV, plus 2 tributaries. Confirms confluence anatomy for Type 2 diagnosis. |
| 160–175 s | Ph5 Elimination Test at SFJ | Probe returns to groin (SFJ). Colour Doppler: large mixed RED+BLUE signal — bidirectional flow during calf squeeze manoeuvre. Finding: Positive elimination test — SFJ is the escape point for Type 1 shunt. | Groin crease, anterior — elimination test (dist=0.06) | SFJ (✓ exp: SFJ) | Look for N1 alongside N2 at SFJ junction level. | Switch to PW Doppler at SFJ to capture bidirectional spectral waveform (antegrade on squeeze, retrograde on release). | CORRECT | YES | YES GSV proximal trunk | N/A | Video 160 s: Probe physically back at groin. Large colour Doppler signal with both red and blue — bidirectional. This is the elimination test by calf compression at SFJ. |
| 175–183 s | Ph5 Elimination Test at SFJ | PW Doppler spectral trace captured at SFJ. Trace shows UPWARD peak (antegrade, calf squeeze) followed by DOWNWARD deflection (retrograde, release). Finding: POSITIVE elimination test. SFJ confirmed as sole escape point for Type 1 shunt. Final diagnosis: Type 1 + 2 combined shunt. | SFJ, PW Doppler capturing elimination waveform (dist=0.07) | SFJ (✓ exp: SFJ) | Move probe slightly deeper to assess N1 and N2 relationship. | Document findings: SFJ incompetent (Type 1), SPJ incompetent (Type 2). Mark anatomy on skin. Plan CHIVA procedure. | WRONG | YES | YES GSV proximal trunk | N/A | Video 180–183 s: Bidirectional PW Doppler trace confirmed. Upward = antegrade on squeeze, downward = retrograde on release. Classic positive elimination test waveform. |

## 4. Guidance Quality Deep-Dive

For each examination step, the full guidance text is shown alongside the expected clinical next action and a pass/fail verdict.

### Step 1: 0–5 s — Phase 1 – SFJ Assessment

Surgeon action: Surgeon places probe at groin crease, anterior surface. B-mode: small anechoic oval (GSV) visible beside larger CFV. Finding: SFJ identified.
Probe position: Groin crease, anterior (dist=0.05)
Region detected: SFJ  (CORRECT)
System guidance: "Look for N1 alongside N2 at SFJ junction level."
Expected next step: Switch Colour Doppler; confirm reflux at SFJ junction (GSV junction with CFV).
Guidance score: PARTIAL (2/3)
Wrong-region detection: CORRECT — flagged
Reason: Probe is scanning posterior anatomy (SPJ) but the protocol requires anterior examination (SFJ). Patient should face the 
Video notes: Video 0 s: B-mode only, probe at inner groin. CFV + GSV visible as adjacent anechoic ovals.

### Step 2: 10–30 s — Phase 1 – SFJ Assessment

Surgeon action: Colour Doppler activated. Large BLUE fill (CFV antegrade), then mixed red/blue at SFJ junction. At ~20 s: large solid RED blob filling GSV — reflux wave on calf-release. Finding: Significant SFJ incompetence confirmed (Type 1 escape point).
Probe position: SFJ, with Colour Doppler active (dist=0.07)
Region detected: SFJ  (CORRECT)
System guidance: "Move probe to upper thigh, posY≈0.10, to assess GSV trunk."
Expected next step: Confirm SFJ junction anatomy on B-mode, then track GSV distally through thigh to map reflux column.
Guidance score: WRONG (1/3)
Video notes: Video 10–30 s: Colour Doppler shows solid red on calf release = retrograde flow = SFJ reflux. Type 1 shunt escape confirmed.

### Step 3: 40–50 s — Phase 2 – GSV Thigh Tracking

Surgeon action: Probe moved to upper thigh (anterior medial). B-mode: large anechoic GSV oval in fascial compartment. Colour Doppler: solid red fill = retrograde flow in GSV trunk. Finding: GSV reflux confirmed in upper thigh (Hunterian canal zone).
Probe position: Upper thigh, anterior (dist=0.30) — Hunterian proximal
Region detected: GSV-THI  (CORRECT)
System guidance: "Move probe to mid-thigh, posY≈0.25, to assess Hunterian junction."
Expected next step: Continue distally along thigh, confirm continuous GSV reflux column down to calf; look for N3 tributary escape point.
Guidance score: PARTIAL (2/3)
Wrong-region detection: CORRECT — flagged
Reason: Probe detected at SFJ (Saphenofemoral Junction) but expected at GSV-THI (Great Saphenous Vein - Thigh).
Video notes: Video 40–50 s: Probe at mid-thigh, solid red Colour Doppler fill = large-calibre retrograde GSV. Type 1 pathway confirmed in thigh.

### Step 4: 50–65 s — Phase 2 – GSV Thigh Tracking

Surgeon action: PW Doppler activated at mid-thigh. Gate placed on GSV. Spectral trace shows downward waveform below baseline (SVD≈16.6 mm, Angle −60°). Finding: Spectral confirmation of retrograde GSV flow at thigh level.
Probe position: Mid-thigh / Hunterian canal, anterior (dist=0.48)
Region detected: GSV-THI  (CORRECT)
System guidance: "Move probe to popliteal fossa, posY≈0.45, to assess SPJ junction."
Expected next step: Move probe further distally along thigh into distal GSV / Dodd zone; continue mapping reflux column toward calf.
Guidance score: WRONG (1/3)
Video notes: Video 55–65 s: PW Doppler trace clearly below baseline = retrograde. SVD 16.6 mm confirms significant GSV trunk diameter. This is the CHIVA RP N2→N1 confirmation step.

### Step 5: 65–80 s — Phase 2 – GSV Thigh Tracking

Surgeon action: Probe moved to distal thigh (Dodd perforator zone). B-mode: single GSV oval still visible in fascial compartment. Finding: GSV reflux column extends to distal thigh.
Probe position: Distal thigh / Dodd zone, anterior (dist=0.65)
Region detected: GSV-THI  (CORRECT)
System guidance: "Move probe to mid-calf, posY≈0.60, to assess GSV-tributary junction."
Expected next step: Move probe to calf, track GSV into fascial compartment; map reflux extent to calf/ankle.
Guidance score: CORRECT (3/3)
Video notes: Video 65–80 s: Probe moves from knee-level back to mid-thigh. Fascia clearly visible, GSV in compartment throughout. Dodd perforator zone is CHIVA Type 2 sub-type boundary.

### Step 6: 80–100 s — Phase 3 – Lower Leg Scan

Surgeon action: Probe moved to upper calf (just below knee), anterior medial. B-mode: small anechoic GSV oval visible in fascial compartment. Finding: GSV continues into calf with same N2 appearance.
Probe position: Upper calf, anterior medial (seg=1, dist=0.15)
Region detected: GSV-CAL  (CORRECT)
System guidance: "Move probe to correct calf position, posY≈0.60, to assess GSV-tributary."
Expected next step: Scan distally along calf to ankle, map full extent of GSV reflux column; look for N3 tributary escape at calf level.
Guidance score: CORRECT (3/3)
Wrong-region detection: CORRECT — flagged
Reason: Probe is scanning posterior anatomy (SPJ) but the protocol requires anterior examination (GSV-CAL). Patient should face 
Video notes: Video 80–100 s: Probe at upper-mid calf. Camera shows probe at shin level. Small single GSV oval on B-mode.

### Step 7: 100–120 s — Phase 3 – Lower Leg Scan

Surgeon action: Probe swept to lower calf / ankle area. B-mode: small GSV oval at ankle level. Multiple vessels visible at popliteal during transit. Finding: GSV reflux maps to ankle, large cluster at knee = popliteal structures.
Probe position: Lower calf / ankle, anterior medial (seg=1, dist=0.75)
Region detected: GSV-CAL  (CORRECT)
System guidance: "Move probe to ankle, posY≈0.85, to assess GSV termination."
Expected next step: Reassess posterior knee (popliteal fossa) to evaluate SPJ and check for Type 2 shunt via SSV.
Guidance score: PARTIAL (2/3)
Video notes: Video 100–120 s: Probe clearly at ankle. Camera shows probe at foot level. GSV very small calibre at this level.

### Step 8: 130–145 s — Phase 4 – SPJ Assessment

Surgeon action: Probe placed at popliteal fossa (posterior knee). B-mode: two large adjacent anechoic ovals — popliteal vein (N1, larger) and SSV (N2, smaller). Colour Doppler: small RED signal at SSV/popliteal interface = reflux. Finding: SPJ incompetence confirmed — SSV reflux into popliteal (Type 2 escape).
Probe position: Popliteal fossa, posterior (dist=0.87, back of thigh)
Region detected: SPJ  (CORRECT)
System guidance: "Move probe to groin crease, posY≈0.06, to assess SFJ junction."
Expected next step: Confirm SSV reflux with PW Doppler at SPJ; then return to SFJ for elimination test.
Guidance score: WRONG (1/3)
Wrong-region detection: CORRECT — flagged
Reason: Probe is scanning anterior anatomy (GSV-CAL) but the protocol requires posterior examination (SPJ). The patient needs to
Video notes: Video 130 s: Two large ovals at popliteal. Red signal at interface. Video 140 s: PW gate positioned, Angle 60°. This is the Type 2 shunt confirmation — SPJ incompetence.

### Step 9: 145–160 s — Phase 4 – SPJ Assessment

Surgeon action: B-mode at SPJ shows 4 anechoic ovals (popliteal vein, SSV, tributaries). PW Doppler gate set at Angle 60°. Finding: Full SPJ anatomy mapped, multiple vessels at junction.
Probe position: Posterior upper calf / SPJ from below (seg=1, dist=0.03)
Region detected: SPJ  (CORRECT)
System guidance: "Adjust probe to anterior thigh, posY≈0.06, to confirm SFJ junction."
Expected next step: Return probe to groin (SFJ). Perform elimination test with calf compression under PW Doppler.
Guidance score: PARTIAL (2/3)
Video notes: Video 150 s: 4 vessels visible at SPJ — popliteal v., SSV, plus 2 tributaries. Confirms confluence anatomy for Type 2 diagnosis.

### Step 10: 160–175 s — Phase 5 – Elimination Test at SFJ

Surgeon action: Probe returns to groin (SFJ). Colour Doppler: large mixed RED+BLUE signal — bidirectional flow during calf squeeze manoeuvre. Finding: Positive elimination test — SFJ is the escape point for Type 1 shunt.
Probe position: Groin crease, anterior — elimination test (dist=0.06)
Region detected: SFJ  (CORRECT)
System guidance: "Look for N1 alongside N2 at SFJ junction level."
Expected next step: Switch to PW Doppler at SFJ to capture bidirectional spectral waveform (antegrade on squeeze, retrograde on release).
Guidance score: CORRECT (3/3)
Video notes: Video 160 s: Probe physically back at groin. Large colour Doppler signal with both red and blue — bidirectional. This is the elimination test by calf compression at SFJ.

### Step 11: 175–183 s — Phase 5 – Elimination Test at SFJ

Surgeon action: PW Doppler spectral trace captured at SFJ. Trace shows UPWARD peak (antegrade, calf squeeze) followed by DOWNWARD deflection (retrograde, release). Finding: POSITIVE elimination test. SFJ confirmed as sole escape point for Type 1 shunt. Final diagnosis: Type 1 + 2 combined shunt.
Probe position: SFJ, PW Doppler capturing elimination waveform (dist=0.07)
Region detected: SFJ  (CORRECT)
System guidance: "Move probe slightly deeper to assess N1 and N2 relationship."
Expected next step: Document findings: SFJ incompetent (Type 1), SPJ incompetent (Type 2). Mark anatomy on skin. Plan CHIVA procedure.
Guidance score: WRONG (1/3)
Video notes: Video 180–183 s: Bidirectional PW Doppler trace confirmed. Upward = antegrade on squeeze, downward = retrograde on release. Classic positive elimination test waveform.

## 5. Key Findings & Recommendations

[PASS] Region localisation: All positions correctly classified.

[PASS] LLM guidance generation: Responses received for all 11 steps.

[WARN] Guidance clinical correctness below 70%: Only 3/11 steps received CORRECT guidance. PARTIAL: 4, WRONG: 4, NO_RESPONSE: 0. The 12-word guidance constraint may be too tight to include all required clinical terms.

[PASS] LLM prompt clinical context: All required terms present in all prompts.

[INFO] SPJ prompt does not mention SSV explicitly: The SPJ prompt says 'Saphenopopliteal Junction' but does not include 'SSV' or 'small saphenous vein'. For Type 2 shunt guidance, the prompt should explicitly mention SSV and its connection to the popliteal vein to guide the LLM correctly.

[INFO] Elimination test context not injected: When the probe returns to SFJ at the end of the examination (elimination test), the system treats it identically to the initial SFJ scan. It does not know the surgeon is performing the elimination test. A phase-aware context field or confirmation of prior SPJ findings would help the guidance explicitly suggest calf compression and bidirectional PW Doppler.