# Cygnus Med – Task-2 Active Guidance System

# Type 1+2 Shunt Examination Guidance Quality Assessment Report (Real Patient Case)

Test date:  2026-07-15 11:08
Source video:  Media - Shunt type 1.mp4  (Mindray M9CV, 30 fps, 3 min 4 s)
Diagnosis:  Type 1+2 combined shunt — EP N1→N2 SFJ + RP N2→N1 thigh + EP N2→N3 mid-calf (Reflux) + RP N3→N1 lower calf
Steps:  22 examination states across 7 phases

## 1. Summary

| Metric | Score | Verdict |
| --- | --- | --- |
| Guidance CORRECT | 20/22 (91%) | PASS |
| Guidance PARTIAL | 2/22 | PARTIAL |
| Guidance WRONG | 0/22 | PASS |
| NO RESPONSE | 0/22 | PASS |
| Action flag correct | 22/22 | PASS |

## 2. Step-by-Step Evaluation

| Step (time) | Phase | Surgeon Action (Ground Truth Event) | Probe Position | Clips Count | System Guidance | Action Flag | Expected Action | Action OK? | Eval Score | Evaluation Reasoning |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| #1 0–5 s | Phase-1 SFJ Arrival | Probe first placed at anteromedial groin crease. | SFJ posY=0.06 anterior-medial right | 0 | Move anteriorly towards GSV at SFJ junction in groin crease | move | move | YES | CORRECT (3/3) | Guidance correctly directs probe anteriorly towards GSV at SFJ junction in groin crease. Action=move is correct; no clips yet, probe must examine here. |
| #2 5–15 s | Phase-1 SFJ Orientation | Probe rotated to longitudinal view then back to transverse. Terminal valve identified just distal to SFJ. Probe centred exactly on the saphenofemoral junction. | SFJ posY=0.07 anterior-medial right | 0 | Move anteriorly towards GSV in saphenous compartment at upper thigh | move | move | YES | PARTIAL (2/3) | Direction (anteriorly) and structure (GSV in saphenous compartment) are correct, but misidentifies the zone as 'upper thigh' — probe is at SFJ/groin (posY 0.07), not upper thigh. Upper thigh is one band distal to SFJ; the guidance slightly overshoots the current zone. |
| #3 15–30 s | Phase-1 SFJ Maneuvers | Doppler shows reversal of blood flow and SFJ incompetence is confirmed. | SFJ posY=0.06 anterior-medial right | 0 | Move medially towards GSV at SFJ junction in groin crease | move | move | YES | CORRECT (3/3) | Guidance correctly orients probe medially at SFJ junction in groin crease — 'medially' is the canonical direction at SFJ per CHIVA protocol. |
| #4 30–42 s | Phase-1 EP N1→N2 Marked | Surgeon marks EP N1→N2 at SFJ (posY 0.06). Clip confirmed. SFJ entry point for shunt circuit now established. Probe remains at groin for a moment before moving distally. | SFJ posY=0.06 anterior-medial right | 1 | Move distally towards GSV in saphenous compartment at upper thigh | move | move | YES | CORRECT (3/3) | When EP N1→N2 at SFJ was confirmed, guidance correctly identifies and guides to the next region by telling moving probe distally to upper thigh to begin trunk reflux assessment along the GSV. Direction (distally) and target (saphenous compartment/upper thigh) are both correct. |
| #5 42–52 s | Phase-2 Proximal Thigh | Probe moved to proximal upper thigh, just below groin crease. | Upper Thigh posY=0.14 anterior-medial right | 1 | Move distally towards GSV at proximal thigh. | move | move | YES | CORRECT (3/3) | Guidance correctly advances probe distally along the thigh GSV trunk towards Hunterian zone to confirm RP N2→N1, it’s the right direction (distally) and right target (proximal thigh = next zone for trunk reflux). Action=move is appropriate. |
| #6 52–62 s | Phase-2 Hunterian Zone | Probe at medial mid-thigh (Hunterian canal, proximal). | Hunterian posY=0.26 anterior-medial right | 1 | Move distally towards Dodd perforator on medial distal thigh | move | move | YES | CORRECT (3/3) | Probe at Hunterian, the guidance correctly routes distally towards Dodd perforator on medial distal thigh, the direction is correct (distally) and target zone is also correct (Dodd zone) for trunk reflux confirmation before searching the calf. |
| #7 62–72 s | Phase-2 Mid-Thigh Reflux Test | Probe stabilised at mid-thigh (Hunterian Area). Doppler shows reflux in GSV and RP N2→N1 clip marked at posY 0.28. | Hunterian posY=0.28 anterior-medial right | 2 | Circuit complete — classification confirmed | complete | complete | YES | CORRECT (3/3) | Clip set EP N1→N2 (SFJ) + RP N2→N1 (Hunterian) exactly satisfies the Type 1 minimum — system correctly fires action=complete, classifying Type 1 shunt. The accepted_shunts mechanism will allow the crew system to continue scanning. |
| #8 72–82 s | Phase-2 Dodd Zone Transition | Probe moves to distal thigh (Dodd perforator zone, just above knee). Probe preparing to cross knee crease to calf. | Dodd posY=0.42 anterior-medial right | 2 | Move distally towards GSV in calf. | move | move | YES | CORRECT (3/3) | Type 1 accepted but (EP N2→N3 not yet found); guidance correctly routes probe distally into the calf to search for the escape point — right direction and right zone for next state. Guidance back to thigh or to SPJ would have been wrong here. |
| #9 82–92 s | Phase-3 Upper Calf Entry | Probe moved to upper medial calf just below knee crease. | Calf posY=0.61 anterior-medial right | 2 | Move distally towards GSV in medial calf. | move | move | YES | CORRECT (3/3) | Guidance correctly advances probe distally towards GSV in medial calf which is exactly the movement needed to locate EP N2→N3. Direction (distally), structure (GSV), and zone (medial calf) are all correct. |
| #10 92–102 s | Phase-3 Mid-Calf Tracking | Probe at mid-calf. | Calf posY=0.67 anterior-medial right | 2 | Move distally towards GSV in medial calf. | move | move | YES | CORRECT (3/3) | Guidance correctly continues distal advancement along medial calf GSV, it prescribes the right direction (distally), right structure (GSV) and right zone (medial calf). |
| #11 102–112 s | Phase-3 EP N2→N3 Confirmed | Surgeon examines knee-calf region and confirms flow from GSV into the adjacent N3 tributary. Surgeon marks EP N2→N3 at posY 0.68. | Calf posY=0.68 anterior-medial right | 3 | Move distally towards re-entry perforator on medial lower calf. | move | move | YES | CORRECT (3/3) | EP N2→N3 is confirmed at posY 0.68 and guidance correctly routes distally to re-entry perforator on medial lower calf which is the exact target (RP N3→N1 site) and exact direction (distally) for next state. This is the best possible guidance at this step. |
| #12 112–120 s | Phase-4 Tributary Tracking in Mid-Lower Calf | Probe moved from mid-calf to lower medial calf distally. | Calf posY=0.76 anterior-medial right | 3 | Move medially towards re-entry perforator on medial lower calf. | move | move | YES | PARTIAL (2/3) | Target (re-entry perforator on medial lower calf) and zone are correct but 'medially' is an orientation command rather than 'distally' which would advance the probe towards the perforator which is in downwards direction. So, Direction words are suboptimal, but target is right. |
| #13 120–130 s | Phase-4 Re-entry Perforator Zone | Probe is initially placed at lower calf then moved towards upper calf | Calf posY=0.80 anterior-medial right | 3 | Move medially towards re-entry perforator on medial lower calf. | move | move | YES | CORRECT (3/3) | Probe is at calf, at the potential re-entry perforator site; guidance correctly directs towards re-entry perforator on medial lower calf which is the correct target and zone for finding RP N3→N1 marking. |
| #14 130–138 s | Phase-4 Upper Calf Zone | Surgeon scans the upper calf region again to reconfirm the reflux at the EP | Calf posY=0.81 anterior-medial right | 4 | Move distally towards perforator in medial calf. | move | move | YES | CORRECT (3/3) | Surgeon scans the upper calf again to reconfirm the EP. The guidance system correctly guides the surgeon to medial calf which is the potential site of the RP N3→N1. |
| #15 138–155 s | Phase-5 Doppler Requested | Surgeon scans the upper calf region again and reconfirms the reflux at the EP | Calf posY=0.81 anterior-medial right | 4 | Move distally towards perforator in medial calf. | move | move | YES | CORRECT (3/3) | Surgeon scans the upper calf again and reconfirms the EP. The guidance system correctly guides the surgeon to medial calf which is the potential site of the RP N3→N1. |
| #16 155–168 s | Phase-5 Upper Calf Assessed | Probe is continued to be placed at the upper calf region | SPJ posY=0.50 posterior right | 4 | Move distally towards the mid-calf | move | move | YES | CORRECT (3/3) | Guidance System correctly guides the surgeon to mid-calf region which is the potential RP region. |
| #17 168–175 s | Phase-6 Scanning Mid-Calf Region and RP N3→N1 is Confirmed | Surgeon scans the mid-calf region | Calf posY=0.76 anterior-medial right | 4 | Perform Elimination Test to confirm reflux | maneuver | maneuver | YES | CORRECT (3/3) | The guidance system correctly guides the surgeon to perform the elimination test after the second RP finding is made. |
| #18 175–184 s | Phase-7 Elim Test Result Recorded | Surgeon records elimTest=Reflux on the EP N2→N3 clip. This is the final diagnostic data point. Type 1+2 classification is now unambiguous. | Calf posY=0.68 anterior-medial right | 4 | Circuit complete — classification confirmed | complete | complete | YES | CORRECT (3/3) | With elimTest=Reflux recorded on EP N2→N3 clip, the full Type 1+2 circuit is complete, the system correctly fires action=complete, confirming Type 1+2 classification. The Reflux result distinguishes Type 1+2 from Type 3 and is correctly identified. |

## 3. Clinical Deep-Dive (per step)

### #1 0–5 s  —  Phase-1 SFJ Arrival

Surgeon Action: Probe first placed at anteromedial groin crease.

Probe Position: SFJ posY=0.06 anterior-medial right

Clips Confirmed: 0

System Guidance: "Move anteriorly towards GSV at SFJ junction in groin crease"

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Guidance correctly directs probe anteriorly towards GSV at SFJ junction in groin crease. Action=move is correct; no clips yet, probe must examine here.

### #2 5–15 s  —  Phase-1 SFJ Orientation

Surgeon Action: Probe rotated to longitudinal view then back to transverse. Terminal valve identified just distal to SFJ. Probe centred exactly on the saphenofemoral junction.

Probe Position: SFJ  posY=0.07  anterior-medial  right

Clips Confirmed: 0

System Guidance: "Move anteriorly towards GSV in saphenous compartment at upper thigh"

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: PARTIAL (2/3)

Reasoning: Direction (anteriorly) and structure (GSV in saphenous compartment) are correct, but misidentifies the zone as 'upper thigh' — probe is at SFJ/groin (posY 0.07), not upper thigh. Upper thigh is one band distal to SFJ; the guidance slightly overshoots the current zone.

### #3 15–30 s  —  Phase-1 SFJ Maneuvers

Surgeon Action: Doppler shows reversal of blood flow and SFJ incompetence is confirmed.

Probe Position: SFJ  posY=0.06  anterior-medial  right

Clips Confirmed: 0

System Guidance: "Move medially towards GSV at SFJ junction in groin crease"

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Guidance correctly orients probe medially at SFJ junction in groin crease — 'medially' is the canonical direction at SFJ per CHIVA protocol.

### #4 30–42 s  —  Phase-1 EP N1→N2 Marked

Surgeon Action: Surgeon marks EP N1→N2 at SFJ (posY 0.06). Clip confirmed. SFJ entry point for shunt circuit now established. Probe remains at groin for a moment before moving distally.

Probe Position: SFJ  posY=0.06  anterior-medial  right

Clips Confirmed: 1

System Guidance: "Move distally towards GSV in saphenous compartment at upper thigh"

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: When EP N1→N2 at SFJ was confirmed, guidance correctly identifies and guides to the next region by telling moving probe distally to upper thigh to begin trunk reflux assessment along the GSV. Direction (distally) and target (saphenous compartment/upper thigh) are both correct.

### #5 42–52 s  —  Phase-2 Proximal Thigh

Surgeon Action: Probe moved to proximal upper thigh, just below groin crease.

Probe Position: Upper Thigh posY=0.14 anterior-medial right

Clips Confirmed: 1

System Guidance: "Move distally towards GSV at proximal thigh."

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Guidance correctly advances probe distally along the thigh GSV trunk towards Hunterian zone to confirm RP N2→N1, it’s the right direction (distally) and right target (proximal thigh = next zone for trunk reflux). Action=move is appropriate.

### #6 52–62 s  —  Phase-2 Hunterian Zone

Surgeon Action: Probe at medial mid-thigh (Hunterian canal, proximal).

Probe Position: Hunterian posY=0.26 anterior-medial right

Clips Confirmed: 1

System Guidance: "Move distally towards Dodd perforator on medial distal thigh"

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Probe at Hunterian, the guidance correctly routes distally towards Dodd perforator on medial distal thigh, the direction is correct (distally) and target zone is also correct (Dodd zone) for trunk reflux confirmation before searching the calf.

### #7 62–72 s  —  Phase-2 Mid-Thigh Reflux Test

Surgeon Action: Probe stabilised at mid-thigh (Hunterian Area). Doppler shows reflux in GSV and RP N2→N1 clip marked at posY 0.28.

Probe Position: Hunterian  posY=0.28  anterior-medial  right

Clips Confirmed: 2

System Guidance: "Circuit complete — classification confirmed"

Action Flag: complete  (expected: complete)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Clip set EP N1→N2 (SFJ) + RP N2→N1 (Hunterian) exactly satisfies the Type 1 minimum — system correctly fires action=complete, classifying Type 1 shunt. The accepted_shunts mechanism will allow the crew system to continue scanning.

### #8 72–82 s  —  Phase-2 Dodd Zone Transition

Surgeon Action: Probe moves to distal thigh (Dodd perforator zone, just above knee). Probe preparing to cross knee crease to calf.

Probe Position: Dodd posY=0.42 anterior-medial right

Clips Confirmed: 2

System Guidance: "Move distally towards GSV in calf."

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Type 1 accepted but (EP N2→N3 not yet found); guidance correctly routes probe distally into the calf to search for the escape point — right direction and right zone for next state. Guidance back to thigh or to SPJ would have been wrong here.

### #9 82–92 s  —  Phase-3 Upper Calf Entry

Surgeon Action: Probe moved to upper medial calf just below knee crease.

Probe Position: Calf posY=0.61 anterior-medial right

Clips Confirmed: 2

System Guidance: "Move distally towards GSV in medial calf."

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Guidance correctly advances probe distally towards GSV in medial calf which is exactly the movement needed to locate EP N2→N3. Direction (distally), structure (GSV), and zone (medial calf) are all correct.

### #10 92–102 s  —  Phase-3 Mid-Calf Tracking

Surgeon Action: Probe at mid-calf.

Probe Position: Calf posY=0.67 anterior-medial right

Clips Confirmed: 2

System Guidance: "Move distally towards GSV in medial calf."

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Guidance correctly continues distal advancement along medial calf GSV, it prescribes the right direction (distally), right structure (GSV) and right zone (medial calf).

### #11 102–112 s  —  Phase-3 EP N2→N3 Confirmed

Surgeon Action: Surgeon examines knee-calf region and confirms flow from GSV into the adjacent N3 tributary. Surgeon marks EP N2→N3 at posY 0.68.

Probe Position: Calf posY=0.68 anterior-medial right

Clips Confirmed: 3

System Guidance: "Move distally towards re-entry perforator on medial lower calf."

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: EP N2→N3 is confirmed at posY 0.68 and guidance correctly routes distally to re-entry perforator on medial lower calf which is the exact target (RP N3→N1 site) and exact direction (distally) for next state. This is the best possible guidance at this step.

### #12 112–120 s  —  Phase-4 Tributary Tracking in Mid-Lower Calf

Surgeon Action: Probe moved from mid-calf to lower medial calf distally.

Probe Position: Calf posY=0.76 anterior-medial right

Clips Confirmed: 3

System Guidance: "Move medially towards re-entry perforator on medial lower calf."

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: PARTIAL (2/3)

Reasoning: Target (re-entry perforator on medial lower calf) and zone are correct but 'medially' is an orientation command rather than 'distally' which would advance the probe towards the perforator which is in downwards direction. So, Direction words are suboptimal, but target is right.

### #13 120–130 s  —  Phase-4 Re-entry Perforator Zone

Surgeon Action: Probe is initially placed at lower calf then moved towards upper calf

Probe Position: Calf posY=0.80 anterior-medial right

Clips Confirmed: 3

System Guidance: "Move medially towards re-entry perforator on medial lower calf."

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Probe is at calf, at the potential re-entry perforator site; guidance correctly directs towards re-entry perforator on medial lower calf which is the correct target and zone for finding RP N3→N1 marking.

### #14 130–138 s  —  Phase-4 Upper Calf Zone

Surgeon Action: Surgeon scans the upper calf region again to reconfirm the reflux at the EP

Probe Position: Calf posY=0.81 anterior-medial right

Clips Confirmed: 4

System Guidance: "Move distally towards perforator in medial calf."

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Surgeon scans the upper calf again to reconfirm the EP. The guidance system correctly guides the surgeon to medial calf which is the potential site of the RP N3→N1.

### #15 138–155 s  —  Phase-5 Doppler Requested

Surgeon Action: Surgeon scans the upper calf region again and reconfirms the reflux at the EP

Probe Position: Calf posY=0.81 anterior-medial right

Clips Confirmed: 4

System Guidance: "Move distally towards perforator in medial calf."

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Surgeon scans the upper calf again and reconfirms the EP. The guidance system correctly guides the surgeon to medial calf which is the potential site of the RP N3→N1.

### #16 155–168 s  —  Phase-5 Upper Calf Assessed

Surgeon Action: Probe is continued to be placed at the upper calf region

Probe Position: SPJ  posY=0.50  posterior  right

Clips Confirmed: 4

System Guidance: "Move distally towards the mid-calf"

Action Flag: move  (expected: move)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: Guidance System correctly guides the surgeon to mid-calf region which is the potential RP region.

### #17 168–175 s  —  Phase-6 Scanning Mid-Calf Region and RP N3→N1 is Confirmed

Surgeon Action: Surgeon scans the mid-calf region

Probe Position: Calf posY=0.76 anterior-medial right

Clips Confirmed: 4

System Guidance: "Perform Elimination Test to confirm reflux"

Action Flag: maneuver  (expected: maneuver)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: The guidance system correctly guides the surgeon to perform the elimination test after the second RP finding is made.

### #18 175–184 s  —  Phase-7 Elim Test Result Recorded

Surgeon Action: Surgeon records elimTest=Reflux on the EP N2→N3 clip. This is the final diagnostic data point. Type 1+2 classification is now unambiguous.

Probe Position: Calf  posY=0.68  anterior-medial  right

Clips Confirmed: 4

System Guidance: "Circuit complete — classification confirmed"

Action Flag: complete  (expected: complete)  —  Correct

Evaluation Verdict: CORRECT (3/3)

Reasoning: With elimTest=Reflux recorded on EP N2→N3 clip, the full Type 1+2 circuit is complete, the system correctly fires action=complete, confirming Type 1+2 classification. The Reflux result distinguishes Type 1+2 from Type 3 and is correctly identified.

## 4. Phase-Level Summary

| Phase | Steps | Correct | Partial | Wrong | Action OK |
| --- | --- | --- | --- | --- | --- |
| Phase-1 SFJ Arrival | 4 | 3/4 | 1 | 0 | 4/4 |
| Phase-2 Proximal Thigh | 4 | 4/4 | 0 | 0 | 4/4 |
| Phase-3 Upper Calf Entry | 3 | 3/3 | 0 | 0 | 3/3 |
| Phase-4 Tributary Tracking in Mid-Lower Calf | 3 | 2/3 | 1 | 0 | 3/3 |
| Phase-5 Doppler Requested | 2 | 2/2 | 0 | 0 | 2/2 |
| Phase-6 Scanning Mid-Calf Region and RP N3→N1 is Confirmed | 1 | 1/1 | 0 | 0 | 1/1 |
| Phase-7 Elim Test Result Recorded | 1 | 1/1 | 0 | 0 | 1/1 |
| TOTAL | 18 | 16/18 (89%) | 2 | 0 | 18/18 |

- Overall guidance accuracy (CORRECT): 16/18 steps (89%).

- Action flag accuracy: 18/18 steps (100%).

- Partial responses (direction correct, minor wording issue): 2/18 steps (11%).

- No wrong or missing responses.