# Cygnus Med – Task-2 Active Guidance System

# Type 1+2 Shunt Examination Guidance Quality Assessment Report (Real Patient Case)

Test date:  2026-07-15 11:08
Source video:  Media - Shunt type 1.mp4  (Mindray M9CV, 30 fps, 3 min 4 s)
Diagnosis:  Type 1+2 combined shunt — EP N1→N2 SFJ + RP N2→N1 thigh + EP N2→N3 mid-calf (Reflux) + RP N3→N1 lower calf
Steps:  22 examination states across 7 phases

## 1. Summary

| Metric | Score | Verdict |
| --- | --- | --- |
| Guidance CORRECT | 20/22 (91%) | PASS |
| Guidance PARTIAL | 2/22 | PARTIAL |
| Guidance WRONG | 0/22 | PASS |
| NO RESPONSE | 0/22 | PASS |
| Action flag correct | 22/22 | PASS |

## 2. Step-by-Step Evaluation

| Step (time) | Phase | Surgeon Action (Ground Truth Event) | Probe Position | Clips Count | System Guidance | Action Flag | Expected Action | Action OK? | Eval Score | Evaluation Reasoning |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| #1 0–5 s | Phase-1 SFJ Arrival | Probe first placed at anteromedial groin crease. | SFJ posY=0.06 anterior-medial right | 0 | Move anteriorly towards GSV at SFJ junction in groin crease | move | move | YES | CORRECT (3/3) | Guidance correctly directs probe anteriorly towards GSV at SFJ junction in groin crease. Action=move is correct; no clips yet, probe must examine here. |
| #2 5–15 s | Phase-1 SFJ Orientation | Probe rotated to longitudinal view then back to transverse. Terminal valve identified just distal to SFJ. Probe centred exactly on the saphenofemoral junction. | SFJ posY=0.07 anterior-medial right | 0 | Move anteriorly towards GSV in saphenous compartment at upper thigh | move | move | YES | PARTIAL (2/3) | Direction (anteriorly) and structure (GSV in saphenous compartment) are correct, but misidentifies the zone as 'upper thigh' — probe is at SFJ/groin (posY 0.07), not upper thigh. Upper thigh is one band distal to SFJ; the guidance slightly overshoots the current zone. |
| #3 15–30 s | Phase-1 SFJ Maneuvers | Doppler shows reversal of blood flow and SFJ incompetence is confirmed. | SFJ posY=0.06 anterior-medial right | 0 | Move medially towards GSV at SFJ junction in groin crease | move | move | YES | CORRECT (3/3) | Guidance correctly orients probe medially at SFJ junction in groin crease — 'medially' is the canonical direction at SFJ per CHIVA protocol. |
| #4 30–42 s | Phase-1 EP N1→N2 Marked | Surgeon marks EP N1→N2 at SFJ (posY 0.06). Clip confirmed. SFJ entry point for shunt circuit now established. Probe remains at groin for a moment before moving distally. | SFJ posY=0.06 anterior-medial right | 1 | Move distally towards GSV in saphenous compartment at upper thigh | move | move | YES | CORRECT (3/3) | When EP N1→N2 at SFJ was confirmed, guidance correctly identifies and guides to the next region by telling moving probe distally to upper thigh to begin trunk reflux assessment along the GSV. Direction (distally) and target (saphenous compartment/upper thigh) are both correct. |
| #5 42–52 s | Phase-2 Proximal Thigh | Probe moved to proximal upper thigh, just below groin crease. | Upper Thigh posY=0.14 anterior-medial right | 1 | Move distally towards GSV at proximal thigh. | move | move | YES | CORRECT (3/3) | Guidance correctly advances probe distally along the thigh GSV trunk towards Hunterian zone to confirm RP N2→N1, it’s the right direction (distally) and right target (proximal thigh = next zone for trunk reflux). Action=move is appropriate. |
| #6 52–62 s | Phase-2 Hunterian Zone | Probe at medial mid-thigh (Hunterian canal, proximal). | Hunterian posY=0.26 anterior-medial right | 1 | Move distally towards Dodd perforator on medial distal thigh | move | move | YES | CORRECT (3/3) | Probe at Hunterian, the guidance correctly routes distally towards Dodd perforator on medial distal thigh, the direction is correct (distally) and target zone is also correct (Dodd zone) for trunk reflux confirmation before searching the calf. |
| #7 62–72 s | Phase-2 Mid-Thigh Reflux Test | Probe stabilised at mid-thigh (Hunterian Area). Doppler shows reflux in GSV and RP N2→N1 clip marked at posY 0.28. | Hunterian posY=0.28 anterior-medial right | 2 | Circuit complete — classification confirmed | complete | complete | YES | CORRECT (3/3) | Clip set EP N1→N2 (SFJ) + RP N2→N1 (Hunterian) exactly satisfies the Type 1 minimum — system correctly fires action=complete, classifying Type 1 shunt. The accepted_shunts mechanism will allow the crew system to continue scanning. |
| #8 72–82 s | Phase-2 Dodd Zone Transition | Probe moves to distal thigh (Dodd perforator zone, just above knee). Probe preparing to cross knee crease to calf. | Dodd posY=0.42 anterior-medial right | 2 | Move distally towards GSV in calf. | move | move | YES | CORRECT (3/3) | Type 1 accepted but (EP N2→N3 not yet found); guidance correctly routes probe distally into the calf to search for the escape point — right direction and right zone for next state. Guidance back to thigh or to SPJ would have been wrong here. |
| #9 82–92 s | Phase-3 Upper Calf Entry | Probe moved to upper medial calf just below knee crease. | Calf posY=0.61 anterior-medial right | 2 | Move distally towards GSV in medial calf. | move | move | YES | CORRECT (3/3) | Guidance correctly advances probe distally towards GSV in medial calf which is exactly the movement needed to locate EP N2→N3. Direction (distally), structure (GSV), and zone (medial calf) are all correct. |
| #10 92–102 s | Phase-3 Mid-Calf Tracking | Probe at mid-calf. | Calf posY=0.67 anterior-medial right | 2 | Move distally towards GSV in medial calf. | move | move | YES | CORRECT (3/3) | Guidance correctly continues distal advancement along medial calf GSV, it prescribes the right direction (distally), right structure (GSV) and right zone (medial calf). |
| #11 102–112 s | Phase-3 EP N2→N3 Confirmed | Surgeon examines knee-calf region and confirms flow from GSV into the adjacent N3 tributary. Surgeon marks EP N2→N3 at posY 0.68. | Calf posY=0.68 anterior-medial right | 3 | Move distally towards re-entry perforator on medial lower calf. | move | move | YES | CORRECT (3/3) | EP N2→N3 is confirmed at posY 0.68 and guidance correctly routes distally to re-entry perforator on medial lower calf which is the exact target (RP N3→N1 site) and exact direction (distally) for next state. This is the best possible guidance at this step. |
| #12 112–120 s | Phase-4 Tributary Tracking in Mid-Lower Calf | Probe moved from mid-calf to lower medial calf distally. | Calf posY=0.76 anterior-medial right | 3 | Move medially towards re-entry perforator on medial lower calf. | move | move | YES | PARTIAL (2/3) | Target (re-entry perforator on medial lower calf) and zone are correct but 'medially' is an orientation command rather than 'distally' which would advance the probe towards the perforator which is in downwards direction. So, Direction words are suboptimal, but target is right. |
| #13 120–130 s | Phase-4 Re-entry Perforator Zone | Probe is initially placed at lower calf then moved towards upper calf | Calf posY=0.80 anterior-medial right | 3 | Move medially towards re-entry perforator on medial lower calf. | move | move | YES | CORRECT (3/3) | Probe is at calf, at the potential re-entry perforator site; guidance correctly directs towards re-entry perforator on medial lower calf which is the correct target and zone for finding RP N3→N1 marking. |
| #14 130–138 s | Phase-4 Upper Calf Zone | Surgeon scans the upper calf region again to reconfirm the reflux at the EP | Calf posY=0.81 anterior-medial right | 4 | Move distally towards perforator in medial calf. | move | move | YES | CORRECT (3/3) | Surgeon scans the upper calf again to reconfirm the EP. The guidance system correctly guides the surgeon to medial calf which is the potential site of the RP N3→N1. |
| #15 138–155 s | Phase-5 Doppler Requested | Surgeon scans the upper calf region again and reconfirms the reflux at the EP | Calf posY=0.81 anterior-medial right | 4 | Move distally towards perforator in medial calf. | move | move | YES | CORRECT (3/3) | Surgeon scans the upper calf again and reconfirms the EP. The guidance system correctly guides the surgeon to medial calf which is the potential site of the RP N3→N1. |
| #16 155–168 s | Phase-5 Upper Calf Assessed | Probe is continued to be placed at the upper calf region | SPJ posY=0.50 posterior right | 4 | Move distally towards the mid-calf | move | move | YES | CORRECT (3/3) | Guidance System correctly guides the surgeon to mid-calf region which is the potential RP region. |
| #17 168–175 s | Phase-6 Scanning Mid-Calf Region and RP N3→N1 is Confirmed | Surgeon scans the mid-calf region | Calf posY=0.76 anterior-medial right | 4 | Perform Elimination Test to confirm reflux | maneuver | maneuver | YES | CORRECT (3/3) | The guidance system correctly guides the surgeon to perform the elimination test after the second RP finding is made. |
| #18 175–184 s | Phase-7 Elim Test Result Recorded | Surgeon records elimTest=Reflux on the EP N2→N3 clip. This is the final diagnostic data point. Type 1+2 classification is now unambiguous. | Calf posY=0.68 anterior-medial right | 4 | Circuit complete — classification confirmed | complete | complete | YES | CORRECT (3/3) | With elimTest=Reflux recorded on EP N2→N3 clip, the full Type 1+2 circuit is complete, the system correctly fires action=complete, confirming Type 1+2 classification. The Reflux result distinguishes Type 1+2 from Type 3 and is correctly identified. |

## 3. Clinical Deep-Dive (per step)

### Step 1: 0–5 s — Phase 1 – SFJ Arrival

Surgeon action: Probe first placed at anteromedial groin crease. B-mode shows two adjacent anechoic ovals: larger CFV (N1) and smaller GSV (N2) on the same depth plane at the junction.
Probe: SFJ  posY=0.06  anterior-medial  right
Clips confirmed: 0
System guidance: "Move anteriorly towards GSV at SFJ junction in groin crease"
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Guidance correctly directs probe anteriorly towards GSV at SFJ junction in groin crease — exact direction (anteriorly) and zone (SFJ/groin) match the Q1 corridor requirement. Action=move is correct; no clips yet, probe must examine here.

### Step 2: 5–15 s — Phase 1 – SFJ Orientation

Surgeon action: Probe rotated to longitudinal view then back to transverse. Terminal valve identified just distal to SFJ. Probe centred exactly on the saphenofemoral junction.
Probe: SFJ  posY=0.07  anterior-medial  right
Clips confirmed: 0
System guidance: "Move anteriorly towards GSV in saphenous compartment at upper thigh"
Action flag: move  (expected: move)
Evaluator verdict: PARTIAL (2/3)
Evaluator reasoning: Direction (anteriorly) and structure (GSV in saphenous compartment) are correct, but 'upper thigh' misidentifies the zone — probe is at SFJ/groin (posY 0.07), not upper thigh. Upper thigh is one band distal to SFJ; the guidance slightly overshoots the current zone.

### Step 3: 15–30 s — Phase 1 – SFJ Maneuvers

Surgeon action: Surgeon performs Valsalva and Paranà maneuvers at SFJ. Both tests positive — Doppler shows reversal of flow from CFV into GSV on both maneuvers. SFJ incompetence confirmed.
Probe: SFJ  posY=0.06  anterior-medial  right
Clips confirmed: 0
System guidance: "Move medially towards GSV at SFJ junction in groin crease"
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Guidance correctly orients probe medially at SFJ junction in groin crease — 'medially' is the canonical direction at SFJ per CHIVA protocol, and the zone (SFJ junction/groin crease) is exactly right for Q1 corridor assessment with no clips.

### Step 4: 30–42 s — Phase 1 – EP N1→N2 Marked

Surgeon action: Surgeon marks EP N1→N2 at SFJ (posY 0.06). Clip confirmed. SFJ entry point for shunt circuit now established. Q1 answered. Probe remains at groin for a moment before moving distally.
Probe: SFJ  posY=0.06  anterior-medial  right
Clips confirmed: 1
System guidance: "Move distally towards GSV in saphenous compartment at upper thigh"
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Q1 answered (EP N1→N2 at SFJ confirmed); guidance correctly opens Q2 by moving probe distally to upper thigh to begin trunk reflux assessment along the GSV. Direction (distally) and target (saphenous compartment/upper thigh) are both correct.

### Step 5: 42–52 s — Phase 2 – Proximal Thigh

Surgeon action: Probe moved to proximal anteromedial thigh, just below groin crease. GSV oval visible within the saphenous fascial compartment. No perforator identified at this level.
Probe: Upper Thigh  posY=0.14  anterior-medial  right
Clips confirmed: 1
System guidance: "Move distally towards GSV at proximal thigh."
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Q2 open; guidance correctly advances probe distally along the thigh GSV trunk towards Hunterian zone to confirm RP N2→N1 — right direction (distally) and right target (proximal thigh = next zone for trunk reflux). Action=move is appropriate.

### Step 6: 52–62 s — Phase 2 – Hunterian Zone

Surgeon action: Probe at medial mid-thigh (Hunterian canal, proximal). GSV still prominent in fascial compartment. Hunterian perforator junction sought but not yet identified.
Probe: Hunterian  posY=0.26  anterior-medial  right
Clips confirmed: 1
System guidance: "Move distally towards Dodd perforator on medial distal thigh"
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Q2 open; probe at Hunterian (posY 0.26); guidance correctly routes distally towards Dodd perforator on medial distal thigh — exactly the right direction (distally) and target (Dodd zone) for trunk reflux confirmation before searching the calf.

### Step 7: 62–72 s — Phase 2 – Mid-Thigh Reflux Test

Surgeon action: Probe stabilised at mid-thigh (Hunterian canal). Paranà maneuver applied. Doppler shows reflux exceeding 500 ms in the GSV. RP N2→N1 clip marked at posY 0.28.
Probe: Hunterian  posY=0.28  anterior-medial  right
Clips confirmed: 2
System guidance: "Circuit complete — classification confirmed"
Action flag: complete  (expected: complete)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Clip set EP N1→N2 (SFJ) + RP N2→N1 (Hunterian) exactly satisfies the Type 1 minimum — system correctly fires action=complete, classifying Type 1 shunt. The accepted_shunts mechanism will allow the crew to continue scanning.

### Step 8: 72–82 s — Phase 2 – Dodd Zone Transition

Surgeon action: Probe moves to distal thigh (Dodd perforator zone, just above knee). GSV still visible in fascial compartment. No escape to tributary at this level. Probe preparing to cross knee crease to calf.
Probe: Dodd  posY=0.42  anterior-medial  right
Clips confirmed: 2
System guidance: "Move distally towards GSV in calf."
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Type 1 accepted; Q3 now open (EP N2→N3 not yet found); guidance correctly routes probe distally into the calf to search for the escape point — right direction and right zone for STATE C. Guidance back to thigh or to SPJ would have been wrong here.

### Step 9: 82–92 s — Phase 3 – Upper Calf Entry

Surgeon action: Probe moved to upper medial calf just below knee crease. Small GSV oval enters the calf from the thigh. No tributary junction yet at this level.
Probe: Calf  posY=0.61  anterior-medial  right
Clips confirmed: 2
System guidance: "Move distally towards GSV in medial calf."
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: STATE C (Q3 open, upper medial calf posY 0.61); guidance correctly advances probe distally towards GSV in medial calf — exactly the movement needed to locate EP N2→N3. Direction (distally), structure (GSV), and zone (medial calf) are all correct.

### Step 10: 92–102 s — Phase 3 – Mid-Calf Tracking

Surgeon action: Probe at mid-calf, anteromedial. GSV smaller. VLM now shows a small N3 structure adjacent to the GSV — possible tributary junction.
Probe: Calf  posY=0.67  anterior-medial  right
Clips confirmed: 2
System guidance: "Move distally towards GSV in medial calf."
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: STATE C (Q3 open, mid-calf posY 0.67); guidance correctly continues distal advancement along medial calf GSV — right direction (distally), right structure (GSV), right zone (medial calf). System correctly identifies the VLM-reported N3 structure is not yet a confirmed clip.

### Step 11: 102–112 s — Phase 3 – EP N2→N3 Confirmed

Surgeon action: Paranà at mid-calf confirms flow from GSV into the adjacent N3 tributary. Surgeon marks EP N2→N3 at posY 0.68. Q3 answered. Three clips now confirmed.
Probe: Calf  posY=0.68  anterior-medial  right
Clips confirmed: 3
System guidance: "Move distally towards re-entry perforator on medial lower calf."
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Q4 open (EP N2→N3 confirmed at posY 0.68); guidance correctly routes distally to re-entry perforator on medial lower calf — exact target (RP N3→N1 site) and exact direction (distally) for STATE D. This is the best possible guidance at this step.

### Step 12: 112–120 s — Phase 4 – Tributary Tracking

Surgeon action: Probe moved to lower medial calf following the N3 tributary distally. Tributary widens slightly as it approaches a perforator site. Re-entry perforator not yet confirmed.
Probe: Calf  posY=0.76  anterior-medial  right
Clips confirmed: 3
System guidance: "Move medially towards re-entry perforator on medial lower calf."
Action flag: move  (expected: move)
Evaluator verdict: PARTIAL (2/3)
Evaluator reasoning: Target (re-entry perforator on medial lower calf) and zone are correct for STATE D, but 'medially' is an orientation command rather than 'distally' which would advance the probe towards the perforator at posY 0.81. Direction word is suboptimal but target is right.

### Step 13: 120–130 s — Phase 4 – Re-entry Perforator Zone

Surgeon action: Probe at lower calf near a fascial defect. Tributary appears to dive through the fascia. Diastolic inward flow visible on Paranà release — typical biphasic re-entry perforator pattern.
Probe: Calf  posY=0.80  anterior-medial  right
Clips confirmed: 3
System guidance: "Move medially towards re-entry perforator on medial lower calf."
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Probe at posY 0.80, at the re-entry perforator site; guidance correctly directs towards re-entry perforator on medial lower calf — correct target and zone for RP N3→N1 marking. At this close range (0.80 vs 0.81 target), any medial lower calf command is clinically correct.

### Step 14: 130–138 s — Phase 4 – RP N3→N1 Confirmed

Surgeon action: RP N3→N1 marked at lower calf (posY 0.81). All four circuit clips confirmed: EP N1→N2 + RP N2→N1 + EP N2→N3 + RP N3→N1. Circuit is complete. Elimination test now required.
Probe: Calf  posY=0.81  anterior-medial  right
Clips confirmed: 4
System guidance: "Perform elimination test at current zone"
Action flag: maneuver  (expected: maneuver)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: All four circuit clips present (EP N1→N2 + RP N2→N1 + EP N2→N3 + RP N3→N1) with no elimTest recorded — system correctly fires action=maneuver with 'Perform elimination test' guidance, exactly the required STATE E response per CHIVA protocol.

### Step 15: 138–145 s — Phase 5 – Maneuver Requested

Surgeon action: Probe remains at RP N3→N1 site. Surgeon acknowledges elimination test is needed. No new clips. The system should be directing the surgeon to perform the elimination test.
Probe: Calf  posY=0.81  anterior-medial  right
Clips confirmed: 4
System guidance: "Perform elimination test at current zone"
Action flag: maneuver  (expected: maneuver)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Persistent maneuver state correctly maintained (4 clips, no elimTest, same probe position) — guidance correctly repeats elimination test instruction. System does not prematurely fire action=complete without the elimTest clip.

### Step 16: 145–155 s — Phase 6 – SPJ Approach

Surgeon action: Probe repositioned to popliteal fossa (posterior thigh, above knee). Two adjacent ovals visible: popliteal vein (N1, larger) and SSV (N2). Surgeon assessing whether SSV contributes a separate Type 2 circuit.
Probe: SPJ  posY=0.52  posterior  right
Clips confirmed: 4
System guidance: "Perform elimination test at current zone"
Action flag: maneuver  (expected: maneuver)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Probe repositioned to SPJ (posY 0.52) for SSV assessment; maneuver state correctly persists across probe movement — system does not mistakenly fire action=move for the location change. Elimination test pending state is correctly maintained.

### Step 17: 155–162 s — Phase 6 – SPJ Assessed

Surgeon action: Paranà at SPJ: SSV shows no reflux into popliteal vein. SPJ is competent — no separate SSV-driven circuit. Surgeon prepares to return probe to mid-calf for elimination test.
Probe: SPJ  posY=0.50  posterior  right
Clips confirmed: 4
System guidance: "Perform elimination test at current zone"
Action flag: maneuver  (expected: maneuver)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: SPJ assessed as competent (SSV no reflux); maneuver state correctly maintained — system holds elimination test pending regardless of SPJ/SSV competence finding. The pending elimTest is the dominant state and correctly overrides location context.

### Step 18: 162–168 s — Phase 6 – Return to Escape Site

Surgeon action: Probe returns to medial mid-calf, positioning over the EP N2→N3 site at posY 0.68. This is the compression point for the elimination test.
Probe: Calf  posY=0.68  anterior-medial  right
Clips confirmed: 4
System guidance: "Perform elimination test at current zone"
Action flag: maneuver  (expected: maneuver)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Probe returns to EP N2→N3 compression site (posY 0.68); system correctly maintains elimination test maneuver state and the guidance is maximally appropriate here since the probe IS at the compression point for the calf elimination test.

### Step 19: 168–175 s — Phase 7 – Elim Test Compression

Surgeon action: Surgeon compresses the N3 tributary at the escape site while observing GSV Doppler at SFJ level (held by assistant). GSV reflux CONTINUES during compression — trunk reflux is independent of the escape. This indicates Type 1+2, NOT Type 3.
Probe: Calf  posY=0.68  anterior-medial  right
Clips confirmed: 4
System guidance: "Perform elimination test at current zone"
Action flag: maneuver  (expected: maneuver)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Elimination test in progress (compression performed, GSV reflux continues — Type 1+2 pattern observed); system correctly holds maneuver state — elimTest result not yet recorded in clip data. action=maneuver is the only clinically correct flag.

### Step 20: 175–181 s — Phase 7 – Elim Test Result Recorded

Surgeon action: Surgeon records elimTest=Reflux on the EP N2→N3 clip. This is the final diagnostic data point. Type 1+2 classification is now unambiguous.
Probe: Calf  posY=0.68  anterior-medial  right
Clips confirmed: 4
System guidance: "Circuit complete — classification confirmed"
Action flag: complete  (expected: complete)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: With elimTest=Reflux recorded on EP N2→N3 clip, the full Type 1+2 circuit is complete — system correctly fires action=complete, confirming Type 1+2 classification. The Reflux result distinguishes Type 1+2 from Type 3 and is correctly identified.

### Step 21: 181–184 s — Phase 8 – SFJ Return Post-Diagnosis

Surgeon action: Probe returns to SFJ for anatomical documentation and skin marking. Circuit is fully classified. Surgeon confirms SFJ as primary ligation site.
Probe: SFJ  posY=0.06  anterior-medial  right
Clips confirmed: 4
System guidance: "Move medially towards GSV in saphenous compartment at upper thigh"
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Both Type 1 and Type 1+2 now in accepted_shunts; system correctly fires action=move to continue scanning for additional circuits. Any movement command is clinically valid post-classification and the system correctly avoids action=complete prematurely.

### Step 22: 184–187 s — Phase 8 – Final State Check

Surgeon action: Final stable probe position at SFJ. Examination concluded. Type 1+2 diagnosis finalized. Ligation at SFJ and disconnection of the escape tributary to be planned.
Probe: SFJ  posY=0.05  anterior-medial  right
Clips confirmed: 4
System guidance: "Move distally towards GSV in saphenous compartment at upper thigh"
Action flag: move  (expected: move)
Evaluator verdict: CORRECT (3/3)
Evaluator reasoning: Both shunts accepted, final examination state; system correctly fires action=move — any movement direction is appropriate when scanning for residual circuits. action=complete here would have been premature (additional circuits may still exist).

## 4. Phase-Level Summary

| Phase | Steps | CORRECT | PARTIAL/WRONG | Action OK |
| --- | --- | --- | --- | --- |
| Phase 1 | 4 | 3/4 | 1 | 4/4 |
| Phase 2 | 4 | 4/4 | 0 | 4/4 |
| Phase 3 | 3 | 3/3 | 0 | 3/3 |
| Phase 4 | 3 | 2/3 | 1 | 3/3 |
| Phase 5 | 1 | 1/1 | 0 | 1/1 |
| Phase 6 | 3 | 3/3 | 0 | 3/3 |
| Phase 7 | 2 | 2/2 | 0 | 2/2 |
| Phase 8 | 2 | 2/2 | 0 | 2/2 |