Manual Testing of Short Scenarios to determine if Shunt System is working well:

Type 1 — SFJ entry + GSV trunk reflux

Important: drag the probe past SPJ zone (posY ≈ 0.50) at some point before marking RP N2→N1 — the system checks that the popliteal region was visited before declaring complete.

Step	Clip	posY	Notes

1	EP N1→N2	0.06	Groin/SFJ

2	—	0.50	Just move probe here (no mark)

3	RP N2→N1	0.25	Mid-thigh

→ Expected: action = complete

→ Result: Scenario Passed and Shunt Identified Correctly

Type 2A — Trunk escape, SFJ competent, no re-entry back to trunk

Step	Clip	posY

1	EP N2→N3	0.30

2	RP N3→N1	0.65

Do NOT mark EP N1→N2 or RP N2→N1.

→ Expected: action = complete

→ Result: Scenario Passed and Shunt Identified Correctly

Type 2B — Perforator feeds GSV, drains via tributary only

Step	Clip	posY

1	EP N2→N2	0.40

2	RP N3→N1	0.65

Do NOT mark EP N1→N2 or RP N2→N1.

→ Expected: action = complete

→ Result: Scenario Passed and Shunt Identified Correctly

Type 2C — Perforator entry + secondary trunk reflux

Step	Clip	posY

1	EP N2→N2	0.40

2	RP N3→N1	0.65

3	RP N2→N1	0.25

Do NOT mark EP N1→N2.

→ Expected: action = complete

→ Result: Scenario Passed and Shunt Identified Correctly

Type 3 — SFJ entry, GSV conduit to escape (no trunk reflux)

Step	Clip	posY

1	EP N1→N2	0.06

2	EP N2→N3	0.25

3	RP N3→N1	0.65

Do NOT mark RP N2→N1.

→ Expected: action = complete

→ Result: Scenario Passed and Shunt Identified Correctly

In elimination test (more clinically rigorous): mark all four clips including RP N2→N1 → system returns action = maneuver → click "No Reflux" in the elimination test banner → action = complete.

Type 1+2 — Same as Type 3 but elimination test shows reflux persists

Step	Clip	posY

1	EP N1→N2	0.06

2	EP N2→N3	0.25

3	RP N3→N1	0.65

4	RP N2→N1	0.40

→ System returns action = maneuver → click "Reflux" in the elimination test banner → action = complete

→ Result: Scenario Passed and Shunt Identified Correctly

Type 4 — Deep blood directly into tributary, trunk refluxes back

Step	Clip	posY

1	EP N1→N3	0.35

2	RP N2→N1	0.25

→ Expected: action = complete (fires immediately — this is the explicit Type 4 complete rule)

→ Result: Scenario Passed and Shunt Identified Correctly

Type 5 — Complex loop: deep→tributary→trunk→tributary→deep

Step	Clip	posY

1	EP N1→N3	0.30

2	RP N3→N2	0.45

3	EP N2→N3	0.55

4	RP N3→N1	0.70

Do NOT mark RP N2→N1.

→ Expected: action = complete

→ Result: Scenario Passed and Shunt Identified Correctly

Type 6 — Pure perforator circuit, trunk completely bypassed

Step	Clip	posY

1	EP N1→N3	0.35

2	RP N3→N1	0.65

Mark absolutely no N2 clips (no EP N1→N2, no EP N2→N3, no RP N2→N1, no RP N3→N2).

→ Expected: action = complete

→ Result: Scenario Passed and Shunt Identified Correctly

Quick reference — few manual mistakes that can happen and how the system would respond to the testing :

| Type | Most common mistake |
| --- | --- |
| 2A/2B | Accidentally adding RP N2→N1 — that changes it to 2C |
| 3 | Adding RP N2→N1 without elimination test — system will demand maneuver |
| 1+2 vs 3 | Same 4 clips; elimination test result is the only differentiator |
| 4 | Adding any RP N3→N1 alongside RP N2→N1 changes it to Type 5 |
| 6 | Any N2 clip (RP N3→N2 or RP N2→N1) would reclassify toward Type 4/5 |

Interface :