# Qwen Local VLM — Thinking vs. Non-Thinking Evaluation (v4)

Model: Qwen3.8-27B (unsloth GGUF, UD-Q4_K_M + mmproj-F16), served via llama.cpp (llama-server) with full GPU offload on an RTX 5090.

Methodology: for each of the 5 task presets, 5 real inputs were run twice, once with the model's reasoning ("thinking") mode enabled, once disabled, for 50 total inference calls. Runs that exhausted their token budget before reaching a final answer are scored as incorrect (no answer given), not excluded. Token budget was set as 3072.

# Summary — All 10 Tables

| Task | Mode | Avg Latency (s) | Accuracy | Sum |
| --- | --- | --- | --- | --- |
| Fascia Depth Based Vein Classification (N1/N2/N3) | Thinking ON | 17.18 s | 100% | 5/5 |
| Fascia Depth Based Vein Classification (N1/N2/N3) | Thinking OFF | 1.47 s | 100% | 5/5 |
| Vein Naming (given real N-class + real blob positions) | Thinking ON | 6.12 s | 80% | 4/5 |
| Vein Naming (given real N-class + real blob positions) | Thinking OFF | 1.69 s | 100% | 5/5 |
| Probe Position: Above vs. At/Below Knee (0/1 binary) | Thinking ON | 11.38 s | 80% | 4/5 |
| Probe Position: Above vs. At/Below Knee (0/1 binary) | Thinking OFF | 2.25 s | 100% | 5/5 |
| Leg Level + Side + Surface (full localisation) | Thinking ON | 27.89 s | 0% | 0/5 |
| Leg Level + Side + Surface (full localisation) | Thinking OFF | 2.49 s | 100% | 5/5 |
| General Free-form Image Q&A (no JSON constraint) | Thinking ON | 14.26 s | 100% | 5/5 |
| General Free-form Image Q&A (no JSON constraint) | Thinking OFF | 8.09 s | 80% | 4/5 |
| ALL TASKS | Thinking ON | 15.36 s | 72% | 18/25 |
| ALL TASKS | Thinking OFF | 3.20 s | 100% | 25/25 |

Findings: (1) The model is strong at both ultrasound tasks — fascia-depth based vein classification scored 100% in BOTH thinking modes (10/10), and vein-naming scored 80% with thinking and 100% without it. (2) Thinking mode's overall accuracy (72%) still is behind non-thinking (96%) almost entirely because of token-budget exhaustion, not reasoning quality: 7 of the 25 thinking-mode runs ran out of budget mid-reasoning and returned no answer at all, heavily concentrated in Task 4 (Leg Level + Side + Surface), which truncated on all 5 runs at its 3072-token cap, this model reasons at real length about facing-direction/toe-vs-heel cues for that specific question, and needs a substantially larger budget to be judged fairly in thinking mode. (3) Non-thinking mode is still much faster overall (3.2s vs. 15.4s average) and ties or wins on 4 of 5 tasks: fascia-depth tied 100/100; vein-naming, probe position, and leg-level were all won outright by non-thinking.

# Task 1 — Fascia Depth Classification (N1/N2/N3)

METHODOLOGY: Uses the simple annotated frames with fascia and veins segmented. Simple prompt having information on how to classify veins based on depth is provided to the LLM.

## Table A — Thinking Enabled

| # | Image | Latency (s) | Correct? | Notes |
| --- | --- | --- | --- | --- |
| 1 | simple_perf_t1.png | 18.59 | Yes | Top blob=N3, bottom blob=N1 — matches ground truth exactly (2/2). |
| 2 | simple_perf_t65.png | 10.85 | Yes | Top-right=N3, near-line blob=N2 — matches ground truth exactly (2/2). |
| 3 | simple_knee_t3.png | 7.01 | Yes | Single blob=N2 — matches ground truth exactly (1/1). |
| 4 | simple_moving_t20.png | 39.15 | Yes | Single blob=N2 — matches ground truth exactly (1/1), though it spent unusually long deliberating over the single-vs-double line ambiguity. |
| 5 | simple_sample_t8.png | 10.31 | Yes | Top-right=N3, near-line blob=N2 — matches ground truth exactly (2/2). |
|  | AVERAGE | 17.18 s | 5/5 | Accuracy = 100% |

## Table B — Thinking Disabled

| # | Image | Latency (s) | Correct? | Notes |
| --- | --- | --- | --- | --- |
| 1 | simple_perf_t1.png | 1.65 | Yes | Top blob=N3, bottom blob=N1 — matches ground truth exactly (2/2). |
| 2 | simple_perf_t65.png | 1.54 | Yes | Top-right=N3, near-line blob=N2 — matches ground truth exactly (2/2). |
| 3 | simple_knee_t3.png | 1.82 | Yes | Single blob=N2 — matches ground truth exactly (1/1). |
| 4 | simple_moving_t20.png | 0.97 | Yes | Single blob=N2 — matches ground truth exactly (1/1). |
| 5 | simple_sample_t8.png | 1.39 | Yes | Top-right=N3, near-line blob=N2 — matches ground truth exactly (2/2). |
|  | AVERAGE | 1.47 s | 5/5 | Accuracy = 100% |

# Task 2 — Vein Naming (given real N-class + real blob positions)

METHODOLOGY: uses the N1/N2/N3 labelled counterparts of the exact same 5 frames used in Task 1, so each blob's N-class is read directly off the image rather than assumed. Probe location (leg_level/surface) is directly provided in context to the LLM and expected names follow directly from the anatomy reference text the model is given, including the explicit one-name-per-frame uniqueness rule for GSV/SSV/CFV/FV/PV.

## Table A — Thinking Enabled

| # | Image | Latency (s) | Correct? | Notes |
| --- | --- | --- | --- | --- |
| 1 | perf_t1.png | 13.80 | No | Exhausted token budget deliberating whether calf N1 should be PV vs. Posterior Tibial/Peroneal Vein; no final answer. |
| 2 | perf_t65.png | 4.03 | Yes | left(N2)=GSV, top-right(N3)=Tributary — correct. |
| 3 | knee_t3.png | 3.81 | Yes | single(N2)=SSV — correct for posterior popliteal fossa. |
| 4 | moving_t20.png | 2.58 | Yes | single=GSV — correct for medial upper thigh. |
| 5 | sample_t8.png | 6.36 | Yes | left(N2)=GSV, top-right(N3)=Tributary — correct. |
|  | AVERAGE | 6.12 s | 4/5 | Accuracy = 80% |

## Table B — Thinking Disabled

| # | Image | Latency (s) | Correct? | Notes |
| --- | --- | --- | --- | --- |
| 1 | perf_t1.png | 1.93 | Yes | top(N3)=Tributary, bottom(N1)=PV — accepted; PV (popliteal vein) is a defensible deep-vein call this close to the popliteal fossa. |
| 2 | perf_t65.png | 1.64 | Yes | left(N2)=GSV, top-right(N3)=Tributary — correct. |
| 3 | knee_t3.png | 1.83 | Yes | single(N2)=SSV — correct. |
| 4 | moving_t20.png | 1.37 | Yes | single=GSV — correct. |
| 5 | sample_t8.png | 1.68 | Yes | left=GSV, top-right(N3)=Tributary — correct. |
|  | AVERAGE | 1.69 s | 5/5 | Accuracy = 100% |

# Task 3 — Probe Position: Above vs. At/Below Knee (0/1binary)

METHODOLOGY:  Any reflux frames were excluded, all 5 below are straightforward scanning frames. The frames wc_0060s/wc_0270s = medial upper thigh, above the knee while frames wc_0750s/wc_0900s/wc_0960s = calf/ankle, below the knee.

## Table A — Thinking Enabled

| # | Image | Latency (s) | Correct? | Notes |
| --- | --- | --- | --- | --- |
| 1 | wc_0060s.jpg | 12.20 | Yes | probe_position=0 (above knee) — correct. |
| 2 | wc_0270s.jpg | 3.33 | Yes | probe_position=0 (above knee) — correct. |
| 3 | wc_0750s.jpg | 11.08 | Yes | probe_position=1 (below knee) — correct. |
| 4 | wc_0900s.jpg | 8.79 | Yes | probe_position=1 (below knee) — correct. |
| 5 | wc_0960s.jpg | 21.50 | No | Exhausted 1536-token budget and there was no final answer. |
|  | AVERAGE | 11.38 s | 4/5 | Accuracy = 80% |

## Table B — Thinking Disabled

| # | Image | Latency (s) | Correct? | Notes |
| --- | --- | --- | --- | --- |
| 1 | wc_0060s.jpg | 2.31 | Yes | probe_position=0 (above knee) — correct. |
| 2 | wc_0270s.jpg | 2.20 | Yes | probe_position=0 (above knee) — correct. |
| 3 | wc_0750s.jpg | 2.27 | Yes | probe_position=1 (below knee) — correct. |
| 4 | wc_0900s.jpg | 2.21 | Yes | probe_position=1 (below knee) — correct. |
| 5 | wc_0960s.jpg | 2.25 | Yes | probe_position=1 (below knee) — correct. |
|  | AVERAGE | 2.25 s | 5/5 | Accuracy = 100% |

# Task 4 — Leg Level + Side + Surface (full localisation)

METHODOLOGY: Same 5 non-reflux frames as Task 3. Scored leniently on the anterior-vs-medial distinction per the task prompt's own instruction that this specific mix-up is non-critical.

## Table A — Thinking Enabled

| # | Image | Latency (s) | Correct? | Notes |
| --- | --- | --- | --- | --- |
| 1 | wc_0060s.jpg | 28.02 | No | Exhausted 3072-token budget on facing-direction deliberation; no final answer. |
| 2 | wc_0270s.jpg | 27.81 | No | Exhausted 3072-token budget; no final answer. |
| 3 | wc_0750s.jpg | 27.84 | No | Exhausted 3072-token budget; no final answer. |
| 4 | wc_0900s.jpg | 27.94 | No | Exhausted 3072-token budget; no final answer. |
| 5 | wc_0960s.jpg | 27.83 | No | Exhausted 3072-token budget; no final answer. |
|  | AVERAGE | 27.89 s | 0/5 | Accuracy = 0% |

## Table B — Thinking Disabled

| # | Image | Latency (s) | Correct? | Notes |
| --- | --- | --- | --- | --- |
| 1 | wc_0060s.jpg | 2.48 | Yes | leg_level=groin_sfj (broadly correct thigh region), surface=anterior (correct group) — correct. |
| 2 | wc_0270s.jpg | 2.41 | Yes | leg_level=groin_sfj, surface=anterior — correct. |
| 3 | wc_0750s.jpg | 2.42 | Yes | leg_level=calf, surface=posterior — correct. |
| 4 | wc_0900s.jpg | 2.63 | Yes | leg_level=calf, surface=posterior — correct. |
| 5 | wc_0960s.jpg | 2.49 | Yes | leg_level=calf, surface=posterior — correct. |
|  | AVERAGE | 2.49 s | 5/5 | Accuracy = 100% |

# Task 5 — General Free-form Image Q&A (no JSON constraint)

METHODOLOGY: Used a sample of images from the ultrasound images set and the webcam images set to ask LLM general questions.

## Table A — Thinking Enabled

| # | Image | Latency (s) | Correct? | Notes |
| --- | --- | --- | --- | --- |
| 1 | knee_t3.png | 15.16 | Yes | Correctly identified the number of veins as 1. |
| 2 | wc_0060s.jpg | 14.56 | Yes | Accurately described the exam-room scene, forearm/smartwatch in foreground. |
| 3 | moving_t20.png | 12.78 | Yes | Correctly described that the position of the vein within the fascial compartment. |
| 4 | wc_0900s.jpg | 14.61 | Yes | Accurately described patient standing, clinician performing ultrasound, assistant present. |
| 5 | wc_0960s.jpg | 14.19 | Yes | Accurately described LA2-14A probe, patient standing on exam table. |
|  | AVERAGE | 14.26 s | 5/5 | Accuracy = 100% |

## Table B — Thinking Disabled

| # | Image | Latency (s) | Correct? | Notes |
| --- | --- | --- | --- | --- |
| 1 | knee_t3.png | 10.06 | No | Correctly identified the number of veins as 1. |
| 2 | wc_0060s.jpg | 8.90 | Yes | Accurately described exam-room scene, arm/smartwatch, clinical setting. |
| 3 | moving_t20.png | 3.61 | Yes | Correctly described that the position of the vein within the fascial compartment. |
| 4 | wc_0900s.jpg | 9.59 | Yes | Accurately described patient standing, doctor performing ultrasound, assistant observing. |
| 5 | wc_0960s.jpg | 8.29 | Yes | Accurately described seated masked examiner, gloves, patient standing. |
|  | AVERAGE | 8.09 s | 5/5 | Accuracy = 100% |

# CONCLUSION

From the results we can conclude that the model with thinking mode disabled is able to provide much quicker as well as accurate results compared to the model with thinking mode enabled which was taking more time to produce results and frequently exceeding the token budget. So it is recommended that we can use the model with thinking mode disabled in the vein annotation and naming tool.