# CHIVA Streaming Guidance — Test Set 2 Report

Date: 2026-06-18  14:03:36    |    API: http://localhost:7861    |    Scenarios: 4

# Overall Summary

| Scenario | Description | Passed | Failed | Timeouts |
| --- | --- | --- | --- | --- |
| type1_extended | Type 1 extended — full leg sweep before confirming RP N2→N1; complete fires at clip_mark when max_visited=0.87. | 23 | 1 | 1 |
| type3_full | Type 3 full — SFJ entry, trunk reflux, Hunterian escape, calf re-entry, maneuver, elimination test (No Reflux), complete. | 21 | 2 | 0 |
| type5_full | Type 5 full — EP N1→N3 → RP N3→N2 → EP N2→N3 → RP N3→N1; no SFJ/SPJ entry, no elimination test, complete at 4th clip. | 21 | 0 | 0 |
| type4_extended | Type 4 extended — bilateral SFJ/SPJ check, full calf perforator sweep, trunk sweep for RP N2→N1; complete at clip_mark (Type 4 rule). | 21 | 0 | 1 |

OVERALL RESULT: FAILURES PRESENT

# Scenario: TYPE1_EXTENDED

Type 1 extended — full leg sweep before confirming RP N2→N1; complete fires at clip_mark when max_visited=0.87.

Steps checked: 24   |   Passed: 23   |   Failed: 1   |   Timeouts: 1

| # | What Surgeon Did | Expected Guidance | LLM Guidance | LLM Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: ts2_type1_ext) | session_ready event | -- | -- | SKIP |
| 2 | Move probe to SFJ posY = 0.06 ¦ anterior-medial surface ¦ right leg | Action = "move" Guidance must mention any of: SFJ, groin, femoral, junction | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 3 | Move probe to SFJ posY = 0.08 ¦ anterior-medial surface ¦ right leg | Action = "move" | -- | -- | TIMEOUT |
| 4 | Mark clip: EP N1→N2 posY = 0.07 ¦ SFJ ¦ right leg | Action = "move" Guidance must mention any of: distal, thigh, GSV, trunk, medial | Move distally along medial thigh toward mid-thigh | move | PASS |
| 5 | Move probe to UPPER_THIGH posY = 0.13 ¦ medial surface ¦ right leg | Action = "move" | Track medial thigh distally following saphenous trunk | move | PASS |
| 6 | Move probe to UPPER_THIGH posY = 0.19 ¦ medial surface ¦ right leg | Action = "move" | Continue down medial thigh toward Hunterian zone | move | PASS |
| 7 | Move probe to HUNTERIAN posY = 0.24 ¦ medial surface ¦ right leg | Action = "move" | Move distally along medial thigh toward mid-thigh | move | PASS |
| 8 | Move probe to HUNTERIAN posY = 0.29 ¦ medial surface ¦ right leg | Action = "move" | Track saphenous trunk distally into the thigh | move | PASS |
| 9 | Move probe to HUNTERIAN posY = 0.34 ¦ medial surface ¦ right leg | Action = "move" | Move distally toward popliteal fossa | move | PASS |
| 10 | Move probe to DISTAL_THIGH posY = 0.40 ¦ medial surface ¦ right leg | Action = "move" | Scan posteriorly at knee level for popliteal junction | move | PASS |
| 11 | Move probe to POPLITEAL posY = 0.47 ¦ posterior surface ¦ right leg | Action = "move" Guidance must mention any of: popliteal, SPJ, posterior, junction | Track medial calf distally following saphenous trunk | move | FAIL |
| x Guidance missing any of: popliteal, SPJ, posterior, junction |  |  |  |  |  |
| 12 | Move probe to POPLITEAL posY = 0.52 ¦ posterior surface ¦ right leg | Action = "move" | Move distally into calf region | move | PASS |
| 13 | Move probe to CALF posY = 0.61 ¦ posterior surface ¦ right leg | Action = "move" Guidance must mention any of: calf, posterior, distal | Follow posterior calf distally toward ankle | move | PASS |
| 14 | Move probe to CALF posY = 0.67 ¦ posterior surface ¦ right leg | Action = "move" | Track medial calf distally following saphenous trunk | move | PASS |
| 15 | Move probe to CALF posY = 0.73 ¦ medial surface ¦ right leg | Action = "move" | Follow medial calf distally toward ankle | move | PASS |
| 16 | Move probe to ANKLE posY = 0.87 ¦ medial surface ¦ right leg | Action = "move" | Scan transversely at medial ankle | move | PASS |
| 17 | Move probe to CALF posY = 0.78 ¦ medial surface ¦ right leg | Action = "move" | Move proximally toward popliteal fossa | move | PASS |
| 18 | Move probe to CALF posY = 0.65 ¦ medial surface ¦ right leg | Action = "move" | Move proximally toward knee level | move | PASS |
| 19 | Move probe to POPLITEAL posY = 0.51 ¦ posterior surface ¦ right leg | Action = "move" | Scan posteriorly at knee level for popliteal junction | move | PASS |
| 20 | Move probe to DISTAL_THIGH posY = 0.41 ¦ medial surface ¦ right leg | Action = "move" | Move distally toward popliteal fossa | move | PASS |
| 21 | Move probe to HUNTERIAN posY = 0.35 ¦ medial surface ¦ right leg | Action = "move" | Scan medially at mid-thigh for Hunterian perforator | move | PASS |
| 22 | Move probe to HUNTERIAN posY = 0.28 ¦ medial surface ¦ right leg | Action = "move" | Move proximally toward groin | move | PASS |
| 23 | Move probe to UPPER_THIGH posY = 0.20 ¦ medial surface ¦ right leg | Action = "move" | Move proximally toward groin crease | move | PASS |
| 24 | Mark clip: RP N2→N1 posY = 0.20 ¦ UPPER_THIGH ¦ right leg | Action = "complete" Guidance must mention any of: circuit, classification, complete, mapped | Circuit mapped — sufficient findings for classification | complete | PASS |
| 25 | Move probe to HUNTERIAN posY = 0.28 ¦ medial surface ¦ right leg | Action = "complete" Guidance must mention any of: circuit, classification | Circuit mapped — sufficient findings for classification | complete | PASS |
| 26 | Move probe to POPLITEAL posY = 0.47 ¦ posterior surface ¦ right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |

# Scenario: TYPE3_FULL

Type 3 full — SFJ entry, trunk reflux, Hunterian escape, calf re-entry, maneuver, elimination test (No Reflux), complete.

Steps checked: 23   |   Passed: 21   |   Failed: 2   |   Timeouts: 0

| # | What Surgeon Did | Expected Guidance | LLM Guidance | LLM Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: ts2_type3_full) | session_ready event | -- | -- | SKIP |
| 2 | Move probe to SFJ posY = 0.06 ¦ anterior-medial surface ¦ right leg | Action = "move" Guidance must mention any of: SFJ, groin, femoral | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 3 | Mark clip: EP N1→N2 posY = 0.06 ¦ SFJ ¦ right leg | Action = "move" Guidance must mention any of: distal, thigh, GSV, trunk | Move distally along medial thigh toward mid-thigh | move | PASS |
| 4 | Move probe to UPPER_THIGH posY = 0.15 ¦ medial surface ¦ right leg | Action = "move" | Track medial thigh distally following saphenous trunk | move | PASS |
| 5 | Mark clip: RP N2→N1 posY = 0.15 ¦ UPPER_THIGH ¦ right leg | Action = "move" Guidance must mention any of: Hunterian, mid-thigh, tributary, escape | Scan medially at mid-thigh for tributary escape | move | PASS |
| 6 | Move probe to HUNTERIAN posY = 0.22 ¦ medial surface ¦ right leg | Action = "move" Guidance must mention any of: Hunterian, tributary, escape, mid-thigh | Scan medially at Hunterian zone for tributary escape | move | PASS |
| 7 | Move probe to HUNTERIAN posY = 0.27 ¦ medial surface ¦ right leg | Action = "move" Guidance must mention any of: Hunterian, tributary, escape, perforator | Scan medially at Hunterian zone for tributary escape | move | PASS |
| 8 | Move probe to HUNTERIAN posY = 0.32 ¦ medial surface ¦ right leg | Action = "move" Guidance must mention any of: Hunterian, tributary, escape | Scan medially at Hunterian zone for tributary escape | move | PASS |
| 9 | Move probe to POPLITEAL posY = 0.47 ¦ posterior surface ¦ right leg | Action = "move" Guidance must mention any of: popliteal, SPJ, posterior | Scan posteriorly at popliteal fossa for SPJ and tributaries | move | PASS |
| 10 | Move probe to POPLITEAL posY = 0.51 ¦ posterior surface ¦ right leg | Action = "move" | Circuit mapped — sufficient findings for classification | complete | FAIL |
| x Action: got "complete", expected "move" |  |  |  |  |  |
| 11 | Move probe to HUNTERIAN posY = 0.29 ¦ medial surface ¦ right leg | Action = "move" Guidance must mention any of: Hunterian, tributary, escape | Circuit mapped — sufficient findings for classification | complete | FAIL |
| x Action: got "complete", expected "move" x Guidance missing any of: Hunterian, tributary, escape |  |  |  |  |  |
| 12 | Mark clip: EP N2→N3 posY = 0.29 ¦ HUNTERIAN ¦ right leg | Action = "move" Must NOT fire action = "maneuver" Guidance must mention any of: tributary, distal, follow | Follow N3 tributary distally on medial thigh | move | PASS |
| 13 | Move probe to DISTAL_THIGH posY = 0.38 ¦ medial surface ¦ right leg | Action = "move" | Follow N3 tributary distally on medial thigh toward knee | move | PASS |
| 14 | Move probe to POPLITEAL posY = 0.44 ¦ medial surface ¦ right leg | Action = "move" | Scan posteriorly at popliteal fossa for re-entry perforator | move | PASS |
| 15 | Move probe to CALF posY = 0.58 ¦ medial surface ¦ right leg | Action = "move" | Follow N3 tributary distally on medial calf | move | PASS |
| 16 | Move probe to CALF posY = 0.64 ¦ medial surface ¦ right leg | Action = "move" | Track medial calf distally for re-entry perforator | move | PASS |
| 17 | Move probe to CALF posY = 0.70 ¦ medial surface ¦ right leg | Action = "move" | Scan medially at calf for re-entry into deep system | move | PASS |
| 18 | Mark clip: RP N3→N1 posY = 0.70 ¦ CALF ¦ right leg | Action = "maneuver" Guidance must mention any of: compress, tributary, Doppler, record | Compress tributary — record whether GSV Doppler changes | maneuver | PASS |
| 19 | Move probe to POPLITEAL posY = 0.46 ¦ posterior surface ¦ right leg | Action = "maneuver" Guidance must mention any of: compress, tributary, record | Compress tributary — record whether GSV Doppler changes | maneuver | PASS |
| 20 | Move probe to DISTAL_THIGH posY = 0.38 ¦ medial surface ¦ right leg | Action = "maneuver" | Compress tributary — record whether GSV Doppler changes | maneuver | PASS |
| 21 | Move probe to HUNTERIAN posY = 0.29 ¦ medial surface ¦ right leg | Action = "maneuver" Guidance must mention any of: compress, tributary, Doppler | Compress tributary — record whether GSV Doppler changes | maneuver | PASS |
| 22 | Mark clip: EP N2→N3 posY = 0.29 ¦ HUNTERIAN ¦ right leg Elimination test: No Reflux | Action = "complete" Guidance must mention any of: circuit, classification, complete, mapped | Circuit mapped — sufficient findings for classification | complete | PASS |
| 23 | Move probe to HUNTERIAN posY = 0.29 ¦ medial surface ¦ right leg | Action = "complete" Guidance must mention any of: circuit, classification | Circuit mapped — sufficient findings for classification | complete | PASS |
| 24 | Move probe to UPPER_THIGH posY = 0.17 ¦ medial surface ¦ right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |

# Scenario: TYPE5_FULL

Type 5 full — EP N1→N3 → RP N3→N2 → EP N2→N3 → RP N3→N1; no SFJ/SPJ entry, no elimination test, complete at 4th clip.

Steps checked: 21   |   Passed: 21   |   Failed: 0   |   Timeouts: 0

| # | What Surgeon Did | Expected Guidance | LLM Guidance | LLM Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: ts2_type5) | session_ready event | -- | -- | SKIP |
| 2 | Move probe to SFJ posY = 0.06 ¦ anterior-medial surface ¦ right leg | Action = "move" Guidance must mention any of: SFJ, groin, femoral | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 3 | Move probe to POPLITEAL posY = 0.47 ¦ posterior surface ¦ right leg | Action = "move" Guidance must mention any of: popliteal, SPJ, posterior | Scan posteriorly at popliteal fossa to locate popliteal junction | move | PASS |
| 4 | Move probe to CALF posY = 0.58 ¦ medial surface ¦ right leg | Action = "move" | Track medial calf distally following saphenous trunk | move | PASS |
| 5 | Move probe to CALF posY = 0.63 ¦ medial surface ¦ right leg | Action = "move" | Move distally along medial calf toward ankle | move | PASS |
| 6 | Move probe to CALF posY = 0.68 ¦ medial surface ¦ right leg | Action = "move" | Follow medial calf distally toward medial ankle | move | PASS |
| 7 | Mark clip: EP N1→N3 posY = 0.68 ¦ CALF ¦ right leg | Action = "move" | Follow N3 tributary distally on medial lower calf | move | PASS |
| 8 | Move probe to CALF posY = 0.63 ¦ medial surface ¦ right leg | Action = "move" | Move distally along medial calf toward site of EP N1→N3 | move | PASS |
| 9 | Move probe to POPLITEAL posY = 0.50 ¦ medial surface ¦ right leg | Action = "move" | Move distally along medial thigh toward calf | move | PASS |
| 10 | Move probe to DISTAL_THIGH posY = 0.41 ¦ medial surface ¦ right leg | Action = "move" | Move distally along medial thigh toward knee | move | PASS |
| 11 | Mark clip: RP N3→N2 posY = 0.41 ¦ DISTAL_THIGH ¦ right leg | Action = "move" | Follow N2 trunk proximally on medial thigh | move | PASS |
| 12 | Move probe to POPLITEAL posY = 0.47 ¦ posterior surface ¦ right leg | Action = "move" | Scan posteriorly at knee level for popliteal junction | move | PASS |
| 13 | Move probe to CALF posY = 0.57 ¦ posterior surface ¦ right leg | Action = "move" | Move proximally along posterior calf toward popliteal fossa | move | PASS |
| 14 | Move probe to CALF posY = 0.63 ¦ posterior surface ¦ right leg | Action = "move" | Move proximally along posterior calf toward popliteal fossa | move | PASS |
| 15 | Mark clip: EP N2→N3 posY = 0.64 ¦ CALF ¦ right leg | Action = "move" Must NOT fire action = "maneuver" | Trace N3 tributary distally toward re-entry perforator | move | PASS |
| 16 | Move probe to CALF posY = 0.70 ¦ posterior surface ¦ right leg | Action = "move" | Move distally along posterior calf toward site of EP N1→N3 | move | PASS |
| 17 | Move probe to CALF posY = 0.74 ¦ lateral surface ¦ right leg | Action = "move" | Move medially across calf to locate site of EP N1→N3 | move | PASS |
| 18 | Move probe to CALF posY = 0.78 ¦ lateral surface ¦ right leg | Action = "move" | Move proximally along lateral calf toward site of EP N1→N3 | move | PASS |
| 19 | Move probe to CALF posY = 0.81 ¦ lateral surface ¦ right leg | Action = "move" | Move medially across calf to locate site of EP N1→N3 | move | PASS |
| 20 | Mark clip: RP N3→N1 posY = 0.81 ¦ CALF ¦ right leg | Action = "complete" Guidance must mention any of: circuit, classification, complete, mapped | Circuit mapped — sufficient findings for classification | complete | PASS |
| 21 | Move probe to CALF posY = 0.81 ¦ lateral surface ¦ right leg | Action = "complete" Guidance must mention any of: circuit, classification | Circuit mapped — sufficient findings for classification | complete | PASS |
| 22 | Move probe to HUNTERIAN posY = 0.30 ¦ medial surface ¦ right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |

# Scenario: TYPE4_EXTENDED

Type 4 extended — bilateral SFJ/SPJ check, full calf perforator sweep, trunk sweep for RP N2→N1; complete at clip_mark (Type 4 rule).

Steps checked: 21   |   Passed: 21   |   Failed: 0   |   Timeouts: 1

| # | What Surgeon Did | Expected Guidance | LLM Guidance | LLM Action | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | Start session (id: ts2_type4_ext) | session_ready event | -- | -- | SKIP |
| 2 | Move probe to SFJ posY = 0.06 ¦ anterior-medial surface ¦ right leg | Action = "move" Guidance must mention any of: SFJ, groin, femoral | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 3 | Move probe to SFJ posY = 0.08 ¦ anterior-medial surface ¦ right leg | Action = "move" | -- | -- | TIMEOUT |
| 4 | Move probe to POPLITEAL posY = 0.47 ¦ posterior surface ¦ right leg | Action = "move" Guidance must mention any of: popliteal, SPJ, posterior | Scan posteriorly at popliteal fossa to locate popliteal junction | move | PASS |
| 5 | Move probe to POPLITEAL posY = 0.52 ¦ posterior surface ¦ right leg | Action = "move" | Scan posteriorly at knee level for popliteal junction | move | PASS |
| 6 | Move probe to CALF posY = 0.58 ¦ medial surface ¦ right leg | Action = "move" | Track medial calf distally following saphenous trunk | move | PASS |
| 7 | Move probe to CALF posY = 0.63 ¦ medial surface ¦ right leg | Action = "move" | Follow N3 tributary distally on medial lower calf | move | PASS |
| 8 | Move probe to CALF posY = 0.68 ¦ medial surface ¦ right leg | Action = "move" | Track medial calf distally following saphenous trunk | move | PASS |
| 9 | Move probe to CALF posY = 0.74 ¦ medial surface ¦ right leg | Action = "move" | Move distally along medial calf toward ankle | move | PASS |
| 10 | Move probe to ANKLE posY = 0.87 ¦ medial surface ¦ right leg | Action = "move" | Scan medially at ankle to locate medial malleolus | move | PASS |
| 11 | Mark clip: EP N1→N3 posY = 0.68 ¦ CALF ¦ right leg | Action = "move" | Follow N3 tributary distally on medial lower calf | move | PASS |
| 12 | Move probe to POPLITEAL posY = 0.48 ¦ posterior surface ¦ right leg | Action = "move" | Move proximally along posterior thigh toward Hunterian zone | move | PASS |
| 13 | Move probe to DISTAL_THIGH posY = 0.40 ¦ medial surface ¦ right leg | Action = "move" | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 14 | Move probe to DISTAL_THIGH posY = 0.37 ¦ medial surface ¦ right leg | Action = "move" | Move proximally along medial thigh toward groin | move | PASS |
| 15 | Move probe to HUNTERIAN posY = 0.33 ¦ medial surface ¦ right leg | Action = "move" | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 16 | Move probe to HUNTERIAN posY = 0.27 ¦ medial surface ¦ right leg | Action = "move" | Move proximally along medial thigh toward groin | move | PASS |
| 17 | Move probe to HUNTERIAN posY = 0.22 ¦ medial surface ¦ right leg | Action = "move" | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 18 | Move probe to UPPER_THIGH posY = 0.17 ¦ medial surface ¦ right leg | Action = "move" | Move proximally along medial thigh toward groin | move | PASS |
| 19 | Move probe to UPPER_THIGH posY = 0.13 ¦ medial surface ¦ right leg | Action = "move" | Scan transversely at groin crease to assess femoral junction | move | PASS |
| 20 | Mark clip: RP N2→N1 posY = 0.30 ¦ HUNTERIAN ¦ right leg | Action = "complete" Guidance must mention any of: circuit, classification, complete, mapped | Circuit mapped — sufficient findings for classification | complete | PASS |
| 21 | Move probe to UPPER_THIGH posY = 0.17 ¦ medial surface ¦ right leg | Action = "complete" Guidance must mention any of: circuit, classification | Circuit mapped — sufficient findings for classification | complete | PASS |
| 22 | Move probe to CALF posY = 0.65 ¦ medial surface ¦ right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |
| 23 | Move probe to POPLITEAL posY = 0.47 ¦ posterior surface ¦ right leg | Action = "complete" | Circuit mapped — sufficient findings for classification | complete | PASS |